ATLAS DISK SECTOR THERMAL QC

| DATE: July 15, 2005 | NAME: Tom/Michael |
| --- | --- |
| SECTOR NUMBER: 9041 |  |
| FRONT VOLTAGE(V): 13.68 | BACK VOLTAGE(V): 13.68 |
| FRONT CURRENT(I): 1.66 | BACK CURRENT(I): 1.66 |
| INLET TEMPERATURE(C): 19.8 | OUTLET TEMPERATURE(C): 20.5 |
| PUMP FLOW RATE(l/min):0.77 | CHILLER SET POINT(C): 18 |
| COMMENTS: IR20050715_0095 Back. Max 30.6 IR20050715_0096 Front Max 33.1 |

FRONT PICTURE

BACK PICTURE