## Summary
| I-V SCORING, LEAKAGE CURRENT EVALUATION OF HYBRIDIZED MODULES | [22 empty] | TOTAL
|  | LAST UPDATED: 13 March 2004, 10:19 | [21 empty] | 352
|  |  | 0 | PASS - NO abnormal behavior (MD or high leakage) observed | [19 empty] | 249 | 70.74%
|  |  | 1 | PASS - Abnormality (MD or high leakage) seen, less than 1 microamp | [19 empty] | 24 | 6.818%
| SCORES |  | 2 | PASS - Abnormality seen, > 1 ua, conditions in 1 hour or less, passes stability test | [19 empty] | 30 | 8.523%
|  |  | 3 | PASS - Abnormality seen, > 1 ua, almost sure chance of conditioning within 24 hours | [19 empty] | 34 | 9.659%
|  |  | 4 | PASS - Abnormality seen, > 1 ua, high leakage current , 1 ua < leakage < 4 ua | [19 empty] | 11 | 3.125%
|  |  | 5 | HOLD - Abnormality seen, > 1 ua, not close to conditioning in 12 hours OR leakage > 4 ua | [19 empty] | 4 | 1.136%
|  |  | 6 | HOLD - Abnormality seen, > 1 ua, conditions in less than 9 hours, but FAILS stability test | [19 empty] | 0 | 0%
|  |  | * | GOLD STAR: No leakage current above 0.5 microamps | [19 empty] | 99 | 28.12%
|  |  | ** | DOUBLE GOLD STAR: No leakage current above 0.4 microamps | [19 empty] | 59 | 16.76%
NOTES:
| TEST PROTOCOL: |  |  | All at ~20C, dry air
|  |  | IV: | Steps to 500V, 1 sec wait. If abnormality detected, measurement repeated with 8 sec wait from 300V
|  |  | MDM: | 500 V for typically 12 hrs. I, T sampled every 5 sec. IV measured immediately following end of MDM run
| IV SHAPES: |  | "normal" | curve bends over,changes less than 30% between 300 and 500 volts
|  |  | "abnormal" | curve continues to rise in almost linear fashion ("ohmic"), or even has increasing slope
| STABILITY TEST |  |  | PASSING = Current changes no more than 10 na/hr for last three hours of MDM run
| SCORES |  | 0 | No microdischarge, I(500) below 1 ua, curve has "normal" shape
|  |  | 1 | Microdischarge, even with 8 sec wait, or "abnormal" curve shape, but total I(500) below 1 ua
|  |  | 2-3 | Conditioning time: Time at which I reaches a minimum before slowly rising.
|  |  | 3 | Tests not run more than 12 hours. Current may still be dropping after 12 hours, slope less than or approaching +/- 10 na/hr
|  |  | 2-3 | Base (minimum) current below 1 ua
|  |  | 4 | High leakage current, 1 ua < current < 4 ua after MDM, typically will have "ohmic" IV curve, passes Stability Test criterion
|  |  | 5 | Current > 4 ua, or if below, unusual behavior (e.g. very noisy)
| PASS/HOLD |  |  | Entry of LT indicates module is awaiting 12-hour test.
| Quality CLASS |  |  | Sensors identified by Nobu as being lower quality. Number of sensors in each Quality Class bin.
| Sum 4 sens: |  |  | Sum of individual sensor IV measurements at 500 V. Prior to 10/03 only best (after MDM) curve saved
After 10/03, if conditioning needed, first meas still saved, summed in this column
| Sum 4 sens: after MDM |  |  | Sum including after-conditioning curves from those sensors needing conditioning. This is the BEST IV sensor sum
| Module afterbonding: |  |  | After 11/03, this is first measurement of hybridized module. If above 1 ua, module is held for 12-hour test
If module was bonded prior to late Nov 03, modules were re-measured to determine IV Score.
| Module Post-MDM |  |  | If module underwent LT (Long Term) test, this is I(500) immediately following test
| Time to baseline |  |  | During MDM run, time to reach minimum current. Note, current typically rises for rest of run.
NA means current never dropped, always rose, so no minimum. (No MD behavior seen.)
| Base current |  |  | Minimum value of current. Note, Post-MDM I(500) is usually .05 ua or so higher
| Plotted |  |  | Data have been analyzed, plotted and collected into Excel summary file.
FILE NAME IS "Pnnn_I-V_data_plots.xls"
| Module ID | [18 empty] | I(500) |  |  | MDM (Microdischarge Decay Measurement)
| P# | Hybrid | Date tested | PASS/HOLD | I-V | [7 empty] | Gold |  |  | Quality CLASS |  |  | Sum 4 sens | Sum 4 sens | Module | Module | Time to | Base | Plotted | Comment
[5 empty] | Score | [7 empty] | Stars |  |  | 2-3 | 4-5 | 6+ |  | (after MDM) | afterbonding | Post-MDM | baseline | current
P | 1 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.508 | [5 empty] | yes
P | 2 | 008 | 2004-01-18 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.481 |  | 2.699 | 1.11 | >12 hrs | 1 | yes | Ohmic IV, I > 1 ua
P | 3 | 010 | 2003-12-07 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.515 |  | 21.3 | 0.391 | 1.5 hrs | 0.37 | yes
P | 4 | 009 | 2003-12-13 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.74 |  | 1.254 | 1.624 | NA | 1.621 | yes | Ohmic IV, I slowly rose over 12 hrs.
P | 5 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.56 | [5 empty] | yes
P | 6 | 011 | 2003-12-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.414 |  | 2.317 | 0.836 | >12 hrs | 0.8 | yes | slope OK
P | 7 | 012 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.511 |  |  |  | yes
P | 8 | 014 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | No data |  | 0.917 |  |  |  | yes
P | 9 | 016 | 2003-12-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.372 |  | 18.608 | 0.393 | 2 hrs | 0.357 | yes
P | 10 | 017 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.332 |  | 0.672 |  |  |  | yes
P | 11 | 018 | 2003-01-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.373 |  | 0.738 |  |  |  | yes | not retested, at CERN
P | 12 | 019 | 2003-12-27 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.386 |  | 5.87 | 0.403 | 8 hrs | 0.405 | yes
P | 13 | 023 | 2003-12-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.386 |  | 1.635 | 0.488 | 6 hrs | 0.48 | yes
P | 14 | 015 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.46 |  | 0.558 |  |  |  | yes
P | 15 | 020 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.416 |  | 0.639 |  |  |  | yes
P | 16 | 037 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.349 |  | 0.463 |  |  |  | yes
P | 17 | 038 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.384 |  | 0.391 |  |  |  | yes
P | 18 | 028 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.397 |  | 0.387 |  |  |  | yes
P | 19 | 040 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.386 |  | 0.355 |  |  |  | yes | 8 sec wait eliminated MD
P | 20 | 024 | 2003-12-04 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.331 |  | 6.66 | 0.351 | 15 min | 0.32 | yes
P | 21 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.332 | [5 empty] | yes
P | 22 | 029 | mech fail |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.353 | [5 empty] | yes
P | 23 | 041 | 2003-02-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.333 |  | 0.706 |  |  |  | yes | not retested, at CERN
P | 24 | 025 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.423 |  | 0.755 |  |  |  | yes
P | 25 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 98.3 | [5 empty] | yes | One really bad sensor
P | 26 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.344 | [5 empty] | yes
P | 27 | 042 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.345 |  | 0.35 |  |  |  | yes
P | 28 | 030 | 2003-12-04 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.378 |  | 21.3 | 2.3 | >12 hrs | 2.5 | yes | Very noisy
P | 29 | 032 | 2003-12-04 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.394 |  | 1.019 |  |  |  | yes
P | 30 | 031 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.478 | [5 empty] | yes | one sensor x2 high
P | 31 | 043 | 2003-03-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.373 |  | 0.5 |  |  |  | yes | not retested, at CERN
P | 32 | 027 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.368 |  | 0.352 |  |  |  | yes
P | 33 | 044 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.46 |  | 0.398 |  |  |  | yes
P | 34 | 069 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.558 |  | 0.485 |  |  |  | yes
P | 35 | 045 | 2003-12-08 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.563 |  | 0.526 |  |  |  | yes
P | 36 | 049 | 2003-12-20 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 2.86 |  | 7.53 | 0.505 | 1 hr | 0.45 | yes
P | 37 | 051 | 2003-12-25 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.588 |  | 2.97 | 0.528 | 1 hr | 0.47 | yes
P | 38 | 033 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.606 |  | 0.442 |  |  |  | yes
P | 39 | 061 | 2003-12-27 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.494 |  | 5.034 | 0.496 | 10 min | 0.44 | yes
P | 40 | 034 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.511 |  | 0.445 |  |  |  | yes
P | 41 | 053 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.435 |  | 0.375 |  |  |  | yes
P | 42 | 036 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.412 |  | 0.373 |  |  |  | yes
P | 43 | 048 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.468 |  | 0.368 |  |  |  | yes
P | 44 | 063 | 2003-12-09 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.497 |  | 2.97 | 0.474 | 12 min | 0.385 | yes
P | 45 | 055 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.496 | [5 empty] | yes
P | 46 | 062 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.462 |  | 0.442 |  |  |  | yes
P | 47 | 065 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.455 |  | 0.507 |  |  |  | yes
P | 48 | 070 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.435 |  | 0.395 |  |  |  | yes
P | 49 | 013 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.447 |  | 0.42 |  |  |  | yes
P | 50 | 072 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.523 |  | 0.52 |  |  |  | yes
P | 51 | 026 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.485 |  | 0.412 |  |  |  | yes
P | 52 | 057 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.345 |  | 0.592 |  |  |  | yes
P | 53 | 071 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.372 |  | 0.417 |  |  |  | yes
P | 54 | 075 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 |  | 0.716 |  |  |  | yes
P | 55 | 074 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.404 |  | 0.482 |  |  |  | yes
P | 56 | 076 | 2003-12-25 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.363 |  | 5.4 | 1.755 | >12 hrs | 1.87 | yes | Ohmic IV, I>1ua
P | 57 | 080 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.409 |  | 0.72 |  |  |  | yes
P | 58 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.422 | [5 empty] | yes | one sensor data missing, entry from logbook
P | 59 | 077 | 2003-12-22 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.356 |  | 0.65 |  |  |  | yes
P | 60 | 078 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.425 |  | 0.37 |  |  |  | yes
P | 61 | 082 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.339 |  | 0.351 |  |  |  | yes
P | 62 | 079 | 2003-12-22 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.392 |  | 6.38 | 0.778 | 2 hrs | 0.77 | yes | bit noisy on base
P | 63 | 081 | 2003-12-22 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.355 |  | 21 | 21 | >12 hrs | 21 ua | yes | 21 ua at 200 V
P | 64 | 083 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.403 |  | 0.386 |  |  |  | yes
P | 65 | 089 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.438 |  | 0.399 |  |  |  | yes
P | 66 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.429 | [5 empty] | yes | 1 sensor slight MD
P | 67 | 095 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.312 |  | 0.257 |  |  |  | yes | 8 sec wait killed MD
P | 68 | 084 | 2003-12-26 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.364 |  | 0.422 |  |  |  | yes
P | 69 | 129 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.52 |  | 1.124 | 0.563 | >12 hrs | 0.55 | yes | slope OK
P | 70 | 090 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.492 |  | 0.388 |  |  |  | yes
P | 71 | 091 | Mech hold |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 |  | 0.496 |  |  |  | yes
P | 72 | 085 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.513 |  | 0.449 |  |  |  | yes
P | 73 | 092 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.457 |  | 0.491 |  |  |  | yes
P | 74 | 149 | 2003-12-26 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.434 |  | 9.849 | 0.407 | 5 min | 0.355 | yes
P | 75 | 086 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.53 |  | 0.443 |  |  |  | yes | 8 sec wait needed
P | 76 | 147 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.769 | [5 empty] | yes | 1 sensor I(500)=.43ua
P | 77 | 088 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.495 |  | 0.413 |  |  |  | yes
P | 78 | 146 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.447 |  | 1.386 | 0.513 | 3.5 hrs | 0.51 | yes | T1 5min; thermistors ng
P | 79 | 073 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.486 |  | 0.447 |  |  |  | yes
P | 80 | 159 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.468 |  | 0.389 |  |  |  | yes
P | 81 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.546 | [5 empty] | yes
P | 82 | 115 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.506 |  | 2.473 | 0.494 | >12 hrs | 0.49 | yes | fast drop 1st hr, slope OK
P | 83 | 144 | 2004-01-06 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.628 |  | 7.61 | 0.482 | >12 hrs | 0.48 | yes | slope OK
P | 84 | 094 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.67 |  | 0.475 |  |  |  | yes
P | 85 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.451 | [5 empty] | yes
P | 86 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.401 | [5 empty] | yes
P | 87 | 093 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.489 |  | 0.402 |  |  |  | yes
P | 88 | 120 | 2003-12-23 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.421 |  | 0.894 |  |  |  | yes
P | 89 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 | [5 empty] | yes
P | 90 | 153 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.428 |  | 0.429 |  |  |  | yes
P | 91 | 178 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 |  | 11.786 | 0.647 | >12 hrs | 0.647 | yes | big drop in 5 min, slope OK
P | 92 | 110 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.469 |  | 0.391 |  |  |  | yes
P | 93 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.465 | [5 empty] | yes
P | 94 | 118 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.501 |  | 0.439 |  |  |  | yes
P | 95 | 123 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.569 |  | 0.42 |  |  |  | yes
P | 96 | 139 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.576 |  | 0.438 |  |  |  | yes
P | 97 | 142 | 2004-01-01 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.484 |  | 1.27 | 0.674 | >12 hrs | 0.65 | yes | slightly irregular decay curve
P | 98 | 087 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.517 |  | 0.427 |  |  |  | yes
P | 99 | 097 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.503 |  | 0.42 |  |  |  | yes
P | 100 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.639 | [5 empty] | yes | 1 sensor I(500) = .30 ua
P | 101 | 121 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.419 |  | 0.323 |  |  |  | yes
P | 102 | 132 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.471 |  | 0.331 |  |  |  | yes
P | 103 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.389 | [5 empty] | yes
P | 104 | 151 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.442 |  | 0.364 |  |  |  | yes
P | 105 | 152 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.406 |  | 0.358 |  |  |  | yes
P | 106 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.467 | [5 empty] | yes
P | 107 | 096 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.455 |  | 0.345 | 0.367 | NA | 0.312 | yes | 12 hr test-normal
P | 108 | 066 | 2004-01-06 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.422 |  | 2.516 | 0.342 | 30 min | 0.305 | yes | rapid drop in 6 min
P | 109 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.468 | [5 empty] | yes
P | 110 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.488 | [5 empty] | yes
P | 111 | 114 | 2004-01-19 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.438 |  | 1.044 | 1.63 | NA | 1.67 | yes | MDM rises
P | 112 | 111 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.523 |  | 0.318 |  |  |  | yes
P | 113 | 067 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.62 |  | 0.372 |  |  |  | yes
P | 114 | 981 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.612 | [5 empty] | yes
P | 115 | 983 | 2004-01-06 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.626 |  | 0.738 |  |  |  | yes | IV ~linear, but <1ua
P | 116 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.716 | [5 empty] | yes | 2 abnormal sensors
P | 117 | 179 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.66 |  | 0.396 |  |  |  | yes
P | 118 | 143 | 2004-01-06 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.612 |  | 1.457 | 0.751 | >12 hrs | 0.75 | yes | slope OK
P | 119 | 060 | 2004-01-06 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.604 |  | 1.209 | 1.499 | NA | 1.55 | yes | No MD, ohmic IV
P | 120 | 113 | 2004-01-06 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 |  | 0.636 |  |  |  | yes | slight MD, pass on 8 sec
P | 121 | 993 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.446 |  | 0.337 |  |  |  | yes
P | 122 | 980 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.463 |  | 0.332 |  |  |  | yes | 8 sec
P | 123 | 105 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.448 |  | 0.336 |  |  |  | yes
P | 124 | 140 | 2003-12-02 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.478 |  | 21.4 | 0.349 | 32 min | 0.32 | yes
P | 125 | 162 | 2003-11-27 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.459 |  | 4.4 | 0.315 | 30 min | 0.3 | yes
P | 126 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | >300 | [5 empty] | yes
P | 127 | 157 | 2003-12-29 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.537 |  | 2.13 | 2.833 | NA | 2.8 | yes | Ohmic IV
P | 128 | 101 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.444 |  | 0.38 |  |  |  | yes
P | 129 | 052 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.439 |  | 0.365 |  |  |  | yes
P | 130 | 164 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.482 |  | 0.406 |  |  |  | yes
P | 131 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | No IV data | [5 empty] | yes
P | 132 | 119 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.462 |  | 0.438 |  |  |  | yes
P | 133 | 184 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.477 |  | 0.419 |  |  |  | yes
P | 134 | 182 | 2004-01-12 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.397 |  |  |  | yes | 8 sec wait killed most MD
P | 135 | 103 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.425 |  | 0.399 |  |  |  | yes
P | 136 | 104 | 2004-01-08 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.428 |  | 7.438 | 0.415 | 5 min | 0.379 | yes
P | 137 | 035 | 2004-01-08 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.43 |  | 10.938 | 2.16 | 5 min | 1.43 | yes | ohmic IV
P | 138 | 154 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.408 |  | 0.411 |  |  |  | yes
P | 139 | 102 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.417 |  | 0.782 |  |  |  | yes
P | 140 | 039 | 2004-01-20 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.491 |  | 6.24 | 0.454 | 1 hr | 0.365 | yes | slope 10 na/hr
P | 141 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.384 | [5 empty] | yes | afterbonding.iv file belongs to P144
P | 142 | 112 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.496 |  | 0.422 |  |  |  | yes | 8 sec IV killed MD
P | 143 | 056 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.364 |  | 0.474 |  |  |  | yes
P | 144 | 130 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.417 |  | 0.349 |  |  |  | yes
P | 145 | 134 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.386 |  | 0.485 |  |  |  | yes
P | 146 | 135 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.54 |  | 0.488 |  |  |  | yes
P | 147 | 116 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.516 |  | 0.534 |  |  |  | yes
P | 148 | 133 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.407 |  | 0.449 |  |  |  | yes
P | 149 | 137 | 2004-01-12 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.393 |  | 0.373 |  |  |  | yes | 8 sec IV killed most MD
P | 150 | 136 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.368 |  | 0.334 |  |  |  | yes
P | 151 | 068 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.414 |  | 0.596 |  |  |  | yes
P | 152 | 141 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.346 |  | 0.322 |  |  |  | yes
P | 153 | 187 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.38 |  | 0.386 |  |  |  | yes
P | 154 | 126 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.359 |  | 0.329 |  |  |  | yes
P | 155 | 138 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 | 1 |  |  | 0.372 |  | 0.33 |  |  |  | yes
P | 156 | 122 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.383 |  | 0.338 |  |  |  | yes
P | 157 | 106 | 2004-01-09 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.347 |  |  |  | yes | 2 sensors w abnormal IV
P | 158 | 191 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.353 |  | 0.331 |  |  |  | yes
P | 159 | 188 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.371 |  | 0.323 |  |  |  | yes
P | 160 | 131 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.31 |  | 0.313 |  |  |  | yes
P | 161 | 193 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.317 |  | 0.315 |  |  |  | yes
P | 162 | 161 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.305 |  | 0.309 |  |  |  | yes
P | 163 | 169 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | inc data |  | 0.285 |  |  |  | yes | missing sensor data, OK
P | 164 | 202 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.331 |  | 0.378 |  |  |  | yes
P | 165 | 203 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.399 |  | 0.648 |  |  |  | yes | ohmic, no MD
P | 166 | 204 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.37 |  | 0.482 |  |  |  | yes
P | 167 | 059 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.405 |  | 0.4 |  |  |  | yes
P | 168 | 195 | Fail-accident |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.563 | 0.532 |  |  |  |  | yes | 1 abnormal sensor
P | 169 | 196 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.46 |  | 0.417 |  |  |  | yes
P | 170 | 198 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.388 |  | 0.399 |  |  |  | yes
P | 171 | 201 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.348 |  | 0.361 |  |  |  | yes
P | 172 | 207 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.435 |  | 0.357 |  |  |  | yes
P | 173 | 199 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 |  | 0.522 |  |  |  | yes
P | 174 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 5.036 | 0.931 |  |  |  |  | yes | One bad sensor, seems to condition
P | 175 | 108 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.385 |  | 0.399 |  |  |  | yes
P | 176 |  | RELEASE |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.759 | 0.403 | [5 empty] | One bad sensor, seems to condition, LOGBOOK 9/15/03 indicates conditioning problems
P | 177 | 180 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.344 |  | 0.433 |  |  |  | yes
P | 178 | 208 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.325 |  | 0.35 |  |  |  | yes
P | 179 | 186 | 2003-12-23 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 8.797 | 0.411 | 6.656 | 0.38 | 1.1 hrs | 0.343 | yes
P | 180 | 209 | mech fail |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.367 | [5 empty] | yes
P | 181 | 226 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.369 |  | 0.352 |  |  |  | yes
P | 182 | 022 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.37 |  | 0.43 |  |  |  | yes
P | 183 | 229 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.391 |  | 0.399 |  |  |  | yes
P | 184 | 232 | chipped TL |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 | [5 empty] | yes
P | 185 | 211 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.448 | 0.393 | 0.401 |  |  |  | yes
P | 186 | 058 | 2003-12-23 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 1.37 | 1.14 | 0.84 |  |  |  | yes
P | 187 | 230 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.398 |  | 0.316 |  |  |  | yes
P | 188 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.355 | [5 empty] | yes
P | 189 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.522 | [5 empty] | yes
P | 190 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.488 | [5 empty] | yes
P | 191 | 214 | 2003-12-23 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.485 |  | 2.639 | 0.383 | 30 min | 0.362 | yes
P | 192 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 385 | [5 empty] | yes | One insanely high sensor
P | 193 | 046 | 2003-12-22 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.472 |  | 0.398 |  |  |  | yes
P | 194 | 054 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.317 |  | 0.293 |  |  |  | yes
P | 195 | 206 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.308 |  | 0.3 |  |  |  | yes
P | 196 | 227 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.37 |  | 0.312 |  |  |  | yes
P | 197 | 166 | 2003-12-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.341 |  | 0.31 |  |  |  | yes
P | 198 | 168 | 2003-12-21 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.312 |  | 1.389 | 0.302 | 15 min | 0.297 | yes
P | 199 | 218 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.377 |  | 0.32 |  |  |  | yes
P | 200 | 216 | PA problem |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.343 | [5 empty] | yes
P | 201 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 73.7 | 0.346 |  |  |  |  | yes | One bad sensor, seems to condition
P | 202 | 219 | 2003-12-23 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.388 |  | 3.439 | 0.358 | 45 min | 0.338 | yes
P | 203 | 220 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 | 1 |  |  | 0.459 |  | 0.348 |  |  |  | yes
P | 204 | 210 | Mech hold |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 2.445 | 0.524 |  |  |  |  | yes | one sensor, high, seems to condition OK
P | 205 | 235 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.431 |  | 0.354 |  |  |  | yes
P | 206 | 237 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.522 |  | 0.386 |  |  |  | yes
P | 207 | 240 | 2004-01-12 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.526 |  | 7.33 | 0.406 | 8 min | 0.36 | yes
P | 208 | 224 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.525 |  | 0.392 |  |  |  | yes
P | 209 | 223 | 2004-01-12 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.572 |  | 10.193 | 0.446 | 25 min | 0.382 | yes
P | 210 | 258 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.491 |  | 0.393 |  |  |  | yes
P | 211 | 250 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.469 |  | 0.389 |  |  |  | yes
P | 212 | 262 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.409 |  | 0.509 |  |  |  | yes
P | 213 | 255 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.492 |  | 0.447 |  |  |  | yes
P | 214 | 213 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.63 |  | 0.465 |  |  |  | yes
P | 215 | 217 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.642 |  | 0.497 |  |  |  | yes
P | 216 | 234 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.642 |  | 0.458 |  |  |  | yes
P | 217 | 236 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.496 |  | 0.437 |  |  |  | yes
P | 218 | 244 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.446 |  | 0.474 |  |  |  | yes
P | 219 | 222 | 2004-01-09 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 13.993 | 0.399 | 10.777 | 0.378 | 8 min | 0.34 | yes
P | 220 | 239 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.544 |  | 0.451 |  |  |  | yes
P | 221 | 242 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.393 |  | 0.357 |  |  |  | yes
P | 222 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.492 | [5 empty] | yes
P | 223 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 347 | 0.445 |  |  |  |  | yes | One chipped sensor, remeas next day was OK??
P | 224 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.539 | 0.474 |  |  |  |  | yes
P | 225 | 241 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.504 |  | 0.388 |  |  |  | yes
P | 226 | 254 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.528 |  | 0.395 |  |  |  | yes
P | 227 | 249 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.552 |  | 0.414 |  |  |  | yes
P | 228 | 263 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.51 |  | 0.415 |  |  |  | yes
P | 229 | 274 | 2004-01-10 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.492 |  | 21 | 0.414 | 3 hrs | 0.395 | yes | good conditioning curve
P | 230 | 173 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.665 |  | 0.561 |  |  |  | yes
P | 231 | 021 | 2004-01-09 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.672 |  | 4.504 | 0.519 | 5 min | 0.442 | yes
P | 232 | 156 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.645 |  | 0.598 |  |  |  | yes
P | 233 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.659 | [5 empty] | yes
P | 234 | 158 | 2004-01-09 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 2.712 | 1.942 | 0.93 |  |  |  | yes | 2 bad sensors
P | 235 | 259 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.475 |  | 0.416 |  |  |  | yes
P | 236 | 267 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.489 |  | 0.424 |  |  |  | yes
P | 237 |  | shipped |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | inc data | [5 empty] | yes | "BURNED"
P | 238 | 270 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.398 |  | 0.394 |  |  |  | yes | 8sec IV killed MD
P | 239 | 266 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.42 |  | 0.455 |  |  |  | yes
P | 240 | 233 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.379 |  | 0.356 |  |  |  | yes
P | 241 | 268 | 2003-12-01 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.467 | 0.385 | 21.4 | 4.99 | >12 hrs | 4.8 | yes | "base" 4.8 ua in 2 hrs
P | 242 | 265 | 2004-01-12 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 21.505 | 0.401 | 5.891 | 0.421 | 40 min | 0.35 | yes
P | 243 | 275 | 2004-01-06 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 1.156 | 0.382 | 5.872 | 0.459 | 8 min | 0.42 | yes
P | 244 | 155 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.38 |  | 0.403 |  |  |  | yes
P | 245 | 145 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.394 |  | 0.422 |  |  |  | yes
P | 246 | 256 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.465 |  | 0.484 |  |  |  | yes
P | 247 | 257 | 2003-11-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.451 |  | 0.458 |  |  |  | yes
P | 248 | 245 | 2003-11-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.456 |  | 0.482 |  |  |  | yes
P | 249 | 176 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.463 |  | 0.453 |  |  |  | yes
P | 250 | 238 | 2003-11-26 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 4 |  |  | 9.5 | 1.85 | 0.628 |  |  |  | yes | 1.8 ua on 1 s wait
P | 251 | 243 | 2003-11-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.462 |  | 0.475 |  |  |  | yes
P | 252 | 287 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.447 |  | 0.477 |  |  |  | yes
P | 253 | 099 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.44 |  | 0.467 |  |  |  | yes
P | 254 |  | no hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.489 | [5 empty] | yes
P | 255 | 289 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.589 |  | 0.586 |  |  |  | yes | early IV 21 ua sens 4
P | 256 | 288 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.613 |  | 0.558 |  |  |  | yes
P | 257 | 299 | 2003-12-04 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 10.5 | 0.736 | 2.79 | 0.999 | >12 hrs | 1 | yes | slope OK
P | 258 | 305 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.612 |  | 0.64 |  |  |  | yes | sensor lkge high
P | 259 | 290 | 2004-01-12 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.649 |  | 1.219 | 1.373 | NA | 1.37 | yes | normal IV, just high
P | 260 | 292 | 2004-01-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.565 |  | 0.592 |  |  |  | yes
P | 261 | 298 | 2003-12-08 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.485 |  | 0.579 |  |  |  | yes | 0.1 ua MD
P | 262 | 260 | 2003-12-10 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.318 |  | 13.81 | 0.33 | 5 min | 0.28 | yes
P | 263 | 160 | 2003-12-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.343 |  | 0.342 |  |  |  | yes
P | 264 | 252 | 2003-12-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.347 |  | 0.35 |  |  |  | yes
P | 265 | 303 | 2003-12-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.339 |  | 0.37 |  |  |  | yes
P | 266 | 253 | 2003-12-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.32 |  | 0.336 |  |  |  | yes
P | 267 | 277 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.485 |  | 0.537 |  |  |  | yes
P | 268 | 272 | 2003-12-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.518 |  | 0.473 |  |  |  | yes
P | 269 | 271 | 2003-12-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.506 |  | 0.479 |  |  |  | yes
P | 270 | 310 | 2003-12-15 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.504 |  | 0.487 |  |  |  | yes
P | 271 |  | no hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.523 | [5 empty] | yes
P | 272 | 280 | 2003-12-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.546 |  | 0.499 |  |  |  | yes
P | 273 | 307 | 2003-12-16 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.504 |  | 11.492 | 0.452 | 5 min | 0.375 | yes
P | 274 | 306 | 2003-12-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.482 |  | 0.398 |  |  |  | yes
P | 275 | 321 | 2003-12-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.487 |  | 0.539 |  |  |  | yes
P | 276 | 313 | 2003-12-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.506 |  | 0.452 |  |  |  | yes
P | 277 | 311 | 2003-12-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.547 |  | 0.51 |  |  |  | yes
P | 278 | 285 | 2003-12-20 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.391 |  | 4.418 | 0.491 | 2 hrs | 0.49 | yes | Bad initial IV ??
P | 279 | 279 | 2003-12-18 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 6.703 | 0.443 | 12.063 | 0.394 | 20 min | 0.342 | yes
P | 280 | 315 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.448 |  | 0.44 |  |  |  | yes
P | 281 | 301 | 2003-12-19 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.431 |  | 0.865 |  |  |  | yes
P | 282 | 304 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.375 |  | 0.398 |  |  |  | yes
P | 283 | 273 | 2004-01-01 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.54 |  | 1.366 | 1.769 | 10 min | 1.46 | yes | erratic, MDM start 21 ua
P | 284 | 317 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.9 |  | 0.921 |  |  |  | yes
P | 285 | 319 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.965 |  | 0.925 |  |  |  | yes
P | 286 | 322 | 2004-01-02 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.867 |  |  |  | yes | ohmic IV curve
P | 287 | 320 | 2004-01-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.392 |  | 1.936 | 0.786 | >12 hrs | 0.78 | yes | slope OK
P | 288 | 221 | 2004-01-07 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.346 |  | 0.353 |  |  |  | yes
P | 289 | 323 | 2004-01-07 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.802 |  | 1.461 | 0.966 | 3 hrs | 0.95 | yes
P | 290 | 324 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.872 |  | 0.957 |  |  |  | yes
P | 291 | 327 | 2004-01-12 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.789 |  | 1.865 | 1.176 | 1.5 hrs | 1.116 | yes | OK, just high
P | 292 | 325 | 2004-01-15 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.764 |  | 0.977 |  |  |  | yes
P | 293 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 23.238 | 17.531 |  |  |  |  | yes | 3 abnormal sensors, Nobu suggested also to hold
P | 294 | 050 | 2004-02-13 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 8.593 | 1.798 | 11.91 | 0.611 | 30 min | 0.586 | yes | One high sensor, but does condition
P | 295 |  | RELEASE |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 3 |  | 30.33 | 22.07 | [5 empty] | One bad sensor, seems to improve with conditioning
P | 296 | 326 | 2004-01-16 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 3 |  | 30.9 | 3.533 | 8.814 | 0.472 | 4 hrs | 0.46 | yes | rapid drop in 1st 6 min
P | 297 |  | release |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 32.66 | 18.46 |  |  |  |  | yes | One bad sensor, seems to improve with conditioning
P | 298 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 3 |  | 39.453 | 21.975 |  |  |  |  | yes | One very bad sensor, Nobu suggested holding
P | 299 | 197 | 2004-01-08 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 4.493 | 0.734 | 1.481 | 0.567 | 1 hr | 0.548 | yes | lgst drop 10 min
P | 300 | 124 | 2004-01-06 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 23.809 | 0.922 | 22.209 | 0.611 | >12 hrs | 0.55 | yes | still shows MD after 12 hrs
P | 301 | 150 | 2004-01-12 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 9.957 | 0.597 | 21 | 0.526 | 4 hrs | 0.5 | yes | 21 ua at 450 V pre MDM
P | 302 | 100 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.491 |  | 0.523 |  |  |  | yes
P | 303 | 098 | 2004-01-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.577 |  | 0.595 |  |  |  | yes
P | 304 | 117 | 2004-01-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.533 |  | 0.567 |  |  |  | yes
P | 305 | 248 | 2004-01-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 0.569 |  | 0.614 |  |  |  | yes
P | 306 | 294 | 2004-01-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.521 |  | 0.489 |  |  |  | yes
P | 307 | 296 | 2004-01-15 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.425 |  | 0.571 |  |  |  | yes
P | 308 | 308 | 2004-01-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.427 |  | 0.472 |  |  |  | yes
P | 309 | 295 | 2004-01-14 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.461 |  | 0.516 |  |  |  | yes
P | 310 | 291 | 2004-01-14 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.445 |  | 0.573 |  |  |  | yes
P | 311 | 312 | 2004-01-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.42 |  | 0.453 |  |  |  | yes
P | 312 | 276 | 2004-01-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 0.545 |  | 0.568 |  |  |  | yes
P | 313 | 269 | 2001-01-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.43 |  | 0.422 |  |  |  | yes | one sensor w unusual IV
P | 314 | 107 | 2004-01-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.326 |  | 0.371 |  |  |  | yes
P | 315 | 309 | 2004-01-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.406 |  | 0.465 |  |  |  | yes
P | 316 |  | Hold |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 1.331 | 1.065 |  |  |  |  | yes | 2 high sensors
P | 317 | 281 | 2004-01-16 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 15.301 | 5.409 | 21.143 | 0.403 | 3 hrs | 0.38 | yes
P | 318 | 247 | 2004-01-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 0.648 | 0.528 | 0.565 |  |  |  | yes
P | 319 | 404 | 2004-03-05 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 10.035 | 5.384 | 5.791 | 0.501 | 24 hrs | 0.5 | yes | Bases in 24 hours
P | 320 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 2 |  | 2.782 | 2.77 |  |  |  |  | yes | One bad sensor, no conditioning effect
P | 321 | 300 | 2004-01-22 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 0.516 |  | 3.574 | 0.49 | 2.5 hrs | 0.465 | yes
P | 322 | 189 | 2004-01-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 2 |  | 0.712 |  | 0.797 |  |  |  | yes
P | 323 | 212 | 2004-01-22 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 2 |  | 0.68 |  | 0.699 |  |  |  | yes | slight MD, w 8sec wait
P | 324 | 174 | 2004-01-23 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 21.624 | 0.607 | 15.75 | 0.524 | 1.2 hrs | 0.475 | yes
P | 325 | 397 | 2004-03-05 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 3.118 | 1.184 | 1.3 | 0.491 | 8 hrs | 0.51 | yes | One bad sensor, seems to respond to conditioning
P | 326 | 148 | 2004-01-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.409 |  | 0.403 |  |  |  | yes
P | 327 | 170 | 2004-01-27 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.347 |  | 0.371 |  |  |  | yes | slight MD, w 8sec wait
P | 328 | 185 | 2004-01-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.38 |  | 0.433 |  |  |  | yes | Bad meas sens 4, extrap sum
P | 329 | 246 | 2004-01-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.414 |  | 0.466 |  |  |  | yes
P | 330 | 175 | 2004-01-23 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 1.638 | 0.952 | 1.429 | 0.636 | >12 hr | 0.62 | yes | one bad sensor,slope OK
P | 331 | 064 | 2004-01-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.446 |  | 0.457 |  |  |  | yes
P | 332 | 177 | 2004-01-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.677 | 0.634 | 0.778 |  |  |  | yes
P | 333 | 215 | 2004-01-27 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.47 |  | 0.525 |  |  |  | yes
P | 334 | 329 | 2004-01-28 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.459 |  | 0.47 |  |  |  | yes
P | 335 | 165 | 2004-01-27 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.389 |  | 0.477 |  |  |  | yes
P | 336 | 109 | 2004-01-27 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.325 |  | 0.387 |  |  |  | yes
P | 337 | 163 | 2004-01-28 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.497 | 0.491 | 0.513 |  |  |  | yes
P | 338 | 047 | 2004-01-28 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.289 |  | 2.543 | 0.801 | >12 hrs | 0.8 | yes | noisy MDM curve
P | 339 |  | Hold |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.339 | [5 empty] | yes
P | 340 | 328 | 2004-01-28 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.377 |  | 0.361 |  |  |  | yes
P | 341 | 293 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.358 |  | 0.373 |  |  |  | yes
P | 342 | 314 | 2004-01-30 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.348 |  | 0.353 |  |  |  | yes
P | 343 | 264 | 2004-01-30 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.339 |  | 0.371 |  |  |  | yes
P | 344 | 375 | 2004-02-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.284 |  | 0.322 |  |  |  | yes
P | 345 | 377 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.31 |  | 0.343 |  |  |  | yes
P | 346 | 379 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.328 |  | 0.354 |  |  |  | yes
P | 347 | 383 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.368 |  | 0.411 |  |  |  | yes
P | 348 | 365 | 2004-02-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.382 |  | 0.37 |  |  |  | yes
P | 349 | 171 | 2004-02-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.471 |  | 0.504 |  |  |  | yes
P | 350 | 357 | 2004-02-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.376 |  | 0.411 |  |  |  | yes
P | 351 | 349 | 2004-02-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.382 |  | 0.384 |  |  |  | yes
P | 352 | 334 | 2004-02-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.369 |  | 0.391 |  |  |  | yes
P | 353 | 333 | 2004-02-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.398 |  | 0.433 |  |  |  | yes
P | 354 | 297 | 2004-02-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.4 |  | 0.432 |  |  |  | yes
P | 355 | 302 | 2004-02-10 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 4.255 | 3.293 | 1.701 | 1.325 | 12 hrs | 1.2 | yes | Acceptable slope in 16 hrs
P | 356 | 355 | 2004-02-10 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 2 |  | 18.635 | 3.027 | 6.739 | 0.349 | 2.5 hrs | 0.349 | yes | very well behaved
P | 357 |  | RELEASE |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 2 | 1 | 17.504 | 6.127 | [5 empty] | One sensor >6ua after conditioning
P | 358 | 354 | 2004-02-10 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.59 | 0.398 | 0.553 |  |  |  | yes | 8 sec wait
P | 359 | 353 | 2004-02-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.348 |  | 0.385 |  |  |  | yes
P | 360 | 382 | 2004-02-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.321 |  | 0.358 |  |  |  | yes
P | 361 | 361 | 2004-02-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.358 |  | 0.385 |  |  |  | yes
P | 362 | 340 | 2004-02-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.35 |  | 0.412 |  |  |  | yes
P | 363 | 337 | 2004-02-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.381 |  | 0.417 |  |  |  | yes
P | 364 | 381 | 2004-02-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.359 |  | 0.374 |  |  |  | yes
P | 365 | 336 | 2004-02-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.428 |  | 0.465 |  |  |  | yes
P | 366 | 378 | 2004-02-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.92 | 0.368 | 0.475 |  |  |  | yes | possibly erroneous hi meas
P | 367 | 251 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.363 |  | 0.42 |  |  |  | yes
P | 368 | 366 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.367 |  | 0.418 |  |  |  | yes
P | 369 | 342 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.344 |  | 0.396 |  |  |  | yes
P | 370 | 332 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.459 |  | 0.512 |  |  |  | yes
P | 371 | 331 | 2004-02-19 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.362 |  | 0.394 |  |  |  | yes
P | 372 | 385 | 2004-02-19 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.354 |  | 0.382 |  |  |  | yes
P | 373 | 363 | 2004-02-19 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  | 1 |  | 0.4 |  | 0.403 |  |  |  | yes
P | 374 | 384 | 2004-02-20 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.402 |  | 0.394 |  |  |  | yes
P | 375 | 231 | 2004-02-20 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.347 |  | 0.328 |  |  |  | yes
P | 376 | 284 | 2004-02-23 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.34 |  | 0.866 |  |  |  | yes
P | 377 | 335 | 2004-02-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.298 |  | 0.309 |  |  |  | yes
P | 378 |  | HOLD |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.376 | 0.343 |  |  |  |  | yes
P | 379 | 338 | 2004-02-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.333 |  | 0.345 |  |  |  | yes
P | 380 |  | HOLD |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.357 | [5 empty] | yes
P | 381 | 362 | 2004-02-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.468 |  | 0.465 |  |  |  | yes
P | 382 | 380 | 2004-02-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.388 |  | 0.388 |  |  |  | yes
P | 383 | 373 | 2004-02-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.485 |  | 0.51 |  |  |  | yes
P | 384 | 356 | 2004-02-25 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 11.243 | 0.607 | 0.766 |  |  |  | yes
P | 385 | 371 | 2004-02-25 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 18.498 | 1.051 | 11.481 | 0.663 | 1.15 hrs | 0.63 | yes | compound MDM slope from different sensors
P | 386 | 330 | 2004-02-25 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 4 |  |  | 0.52 |  | 0.625 |  |  |  | yes
P | 387 | 283 | 2004-02-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 2 |  |  | 4.809 | 0.703 | 11.915 | 0.679 | 1.8 hrs | 0.625 | yes | Long decay from bonding, not sensors
P | 388 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 2 | 1 | 48.378 | 26.413 |  |  |  |  | yes | May possibly condition
P | 389 | 396 | 2004-02-28 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 16.825 | 0.487 | 21.22 | 0.451 | 25 min | 0.41 | yes | Conditions very well
P | 390 | 376 | 2004-02-27 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 5.237 | 0.424 | 11.566 | 0.373 | 30 min | 0.341 | yes | Conditions very well
P | 391 | 359 | 2004-02-27 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.411 |  | 21.4 | 21.1 | >12 hrs | 21 | yes | 21 ua at 200V, at 450 V after 12 hrs
P | 392 | 374 | 2004-03-03 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.482 |  | 0.602 |  |  |  | yes | 1 abnormal sensor, afterbonding IV slightly irregular
P | 393 | 403 | 2004-03-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.666 | 3.14 | 3.485 | 0.62 | 12 hrs | 0.63 | yes | OK to bond
P | 394 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.566 | 0.493 |  |  |  |  | yes | Very noisy, NOT probe-related
P | 395 | 395 | 2004-03-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.596 | 0.481 | 0.608 |  |  |  | yes
P | 396 | 339 | 2004-03-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.457 |  | 0.468 |  |  |  | yes
P | 397 | 402 | 2004-03-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.47 |  | 0.506 |  |  |  | yes
P | 398 | 405 | 2004-03-05 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 13.947 | 0.599 | 6.763 | 0.456 | 15 min | 0.404 | yes | Very fine conditioning
P | 399 | 399 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.438 |  | 0.486 |  |  |  | yes
P | 400 | 400 | 2004-03-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.438 |  | 0.472 |  |  |  | yes
P | 401 | 401 | 2004-03-08 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.369 |  | 1.814 | 0.963 | 12 hrs | 0.963 | yes | ohmic IV curve, noisy MDM
P | 402 | 408 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.374 |  | 0.427 |  |  |  | yes
P | 403 | 416 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.392 |  | 0.379 |  |  |  | yes
P | 404 | 414 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.468 |  | 0.539 |  |  |  | yes
P | 405 | 415 | 2004-03-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.434 |  | 0.472 |  |  |  | yes
P | 406 | 413 | 2004-03-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.427 |  | 0.489 |  |  |  | yes
P | 407 | 419 | 2004-03-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.449 |  | 0.496 |  |  |  | yes
P | 408 | 417 | 2004-03-11 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 9.325 | 0.482 | 15.858 | 0.459 | 45 min | 0.42 | yes | noisy MDM curve
P | 409 | 418 | 2004-03-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 |  | 0.532 |  |  |  | yes
P | 410 | 427 | 2004-03-11 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.581 | 6.631 | 12.05 | 0.504 | 25 min | 0.502 | yes | break @ 300, but does condition
P | 411 | 420 | 2004-03-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.465 |  | 0.549 |  |  |  | yes
P | 412 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.511
P | 413 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.514
P | 414 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.668 | 0.599 | [5 empty] | one sensor somewhat erratic
P | 415 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.404
P | 416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.384
P | 417 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.423
P | 418 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.417
P | 419 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.476
P | 420 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.465
P | 421 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.422
P | 422 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.431
P | 423 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.454
P | 424 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.476
P | 425 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.572
P | 426 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.469
P | 427 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.399
P | 428 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.384
P | 429 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.419
P | 430 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.463
P | 431 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.493
P | 432 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 13.184 | 0.133 | [5 empty] | 1 sensor conditions
P | 433 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.403
P | 434 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.378
P | 435 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.412
P | 436 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.446
P | 437 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.463
P | 438 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.451
P | 439 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.415
P | 440 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.421
P | 441 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.44
P | 442 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.505 | [6 empty] | 1 abnormal sensor
P | 443 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.832 | 0.762 | [5 empty] | 1 abnormal sensor
P | 444 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.426
P | 445 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.436 | [6 empty] | Backside sensors both noisy
P | 446 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.512
P | 447 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.536 | [6 empty] | 1 abnormal sensor
P | 448 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.461
P | 449 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 2.477 | 0.515 | [5 empty] | 2 abnormal sensors, condition OK
P | 450 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.996 | 2.035 | [5 empty] | 2 abnormal sensors, probably will condition
P | 451 |  | IV & glue | hold iv |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.854 | 2.91 | [5 empty] | One very abnormal sensor
P | 452 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 11.859 | 0.54 | [5 empty] | 1 sens, conditions in 1 min
P | 453 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.776 | 0.599 | [5 empty] | 1 abnormal sensor, conditions ~OK
P | 454 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.508
P | 455 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.503
P | 456 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.463
P | 457 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.464
P | 458 |  |  | HOLD |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.96 | 21.91 | [5 empty] | 1 sensor, 21ua @ 30V no condition effect
P | 459 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.377
P | 460 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.467
P | 461 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.433
P | 462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.428
P | 463 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 11.905 | 0.416 | [5 empty] | 1 sensor conditions 2.5 min
P | 464 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 465 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 466 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 467 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 468 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 469 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 470 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 471 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 472 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 473 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 474 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 475 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 477 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 478 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 479 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 480 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 481 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 482 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 483 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 484 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 485 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 486 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 487 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 488 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 489 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 490 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 491 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 492 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 493 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 494 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 495 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 496 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 497 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 498 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 499 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 500 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 501 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 502 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 503 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 504 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 505 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 506 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 507 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 508 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 509 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 510 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 511 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 512 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 513 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 514 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 515 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 516 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 517 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 518 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 519 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 520 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 521 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 522 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 523 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 524 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 525 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 526 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 527 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 528 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 529 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 530 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 531 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 532 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 533 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 534 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 535 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 536 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 537 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 538 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 539 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 540 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 541 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 542 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 543 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 544 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 545 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 546 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 547 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 548 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 549 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 550 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 125 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 127 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 128 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 167 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 172 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 181 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 183 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 190 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 192 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 194 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 200 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 205 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 225 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 228 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 261 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 278 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 282 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 286 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 316 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 318 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
341
343
344
345
346
347
348
350
351
352
358
360
364
367
368
369
370
372
386
387
388
389
390
391
392
393
394
398
406
407
409
410
411
412
421
422
423
424
425
426
428
429
430

## Summary sorted by hybrid
| I-V SCORING, LEAKAGE CURRENT EVALUATION OF HYBRIDIZED MODULES | [22 empty] | TOTAL
|  | LAST UPDATED: 13 March 2004, 10:19 | [21 empty] | 351
|  |  | 0 | PASS - NO abnormal behavior (MD or high leakage) observed | [19 empty] | 249 | 70.94%
|  |  | 1 | PASS - Abnormality (MD or high leakage) seen, less than 1 microamp | [19 empty] | 24 | 6.838%
| SCORES |  | 2 | PASS - Abnormality seen, > 1 ua, conditions in 1 hour or less, passes stability test | [19 empty] | 30 | 8.547%
|  |  | 3 | PASS - Abnormality seen, > 1 ua, almost sure chance of conditioning within 24 hours | [19 empty] | 34 | 9.687%
|  |  | 4 | PASS - Abnormality seen, > 1 ua, high leakage current , 1 ua < leakage < 4 ua | [19 empty] | 10 | 2.849%
|  |  | 5 | HOLD - Abnormality seen, > 1 ua, not close to conditioning in 12 hours OR leakage > 4 ua | [19 empty] | 4 | 1.14%
|  |  | 6 | HOLD - Abnormality seen, > 1 ua, conditions in less than 9 hours, but FAILS stability test | [19 empty] | 0 | 0%
|  |  | * | GOLD STAR: No leakage current above 0.5 microamps | [19 empty] | 99 | 28.21%
|  |  | ** | DOUBLE GOLD STAR: No leakage current above 0.4 microamps | [19 empty] | 59 | 16.81%
NOTES:
| TEST PROTOCOL: |  |  | All at ~20C, dry air
|  |  | IV: | Steps to 500V, 1 sec wait. If abnormality detected, measurement repeated with 8 sec wait from 300V
|  |  | MDM: | 500 V for typically 12 hrs. I, T sampled every 5 sec. IV measured immediately following end of MDM run
| IV SHAPES: |  | "normal" | curve bends over,changes less than 30% between 300 and 500 volts
|  |  | "abnormal" | curve continues to rise in almost linear fashion ("ohmic"), or even has increasing slope
| STABILITY TEST |  |  | PASSING = Current changes no more than 10 na/hr for last three hours of MDM run
| SCORES |  | 0 | No microdischarge, I(500) below 1 ua, curve has "normal" shape
|  |  | 1 | Microdischarge, even with 8 sec wait, or "abnormal" curve shape, but total I(500) below 1 ua
|  |  | 2-3 | Conditioning time: Time at which I reaches a minimum before slowly rising.
|  |  | 3 | Tests not run more than 12 hours. Current may still be dropping after 12 hours, slope less than or approaching +/- 10 na/hr
|  |  | 2-3 | Base (minimum) current below 1 ua
|  |  | 4 | High leakage current, 1 ua < current < 4 ua after MDM, typically will have "ohmic" IV curve, passes Stability Test criterion
|  |  | 5 | Current > 4 ua, or if below, unusual behavior (e.g. very noisy)
| PASS/HOLD |  |  | Entry of LT indicates module is awaiting 12-hour test.
| Quality CLASS |  |  | Sensors identified by Nobu as being lower quality. Number of sensors in each Quality Class bin.
| Sum 4 sens: |  |  | Sum of individual sensor IV measurements at 500 V. Prior to 10/03 only best (after MDM) curve saved
After 10/03, if conditioning needed, first meas still saved, summed in this column
| Sum 4 sens: after MDM |  |  | Sum including after-conditioning curves from those sensors needing conditioning. This is the BEST IV sensor sum
| Module afterbonding: |  |  | After 11/03, this is first measurement of hybridized module. If above 1 ua, module is held for 12-hour test
If module was bonded prior to late Nov 03, modules were re-measured to determine IV Score.
| Module Post-MDM |  |  | If module underwent LT (Long Term) test, this is I(500) immediately following test
| Time to baseline |  |  | During MDM run, time to reach minimum current. Note, current typically rises for rest of run.
NA means current never dropped, always rose, so no minimum. (No MD behavior seen.)
| Base current |  |  | Minimum value of current. Note, Post-MDM I(500) is usually .05 ua or so higher
| Plotted |  |  | Data have been analyzed, plotted and collected into Excel summary file.
FILE NAME IS "Pnnn_I-V_data_plots.xls"
| Module ID | [18 empty] | I(500) |  |  | MDM (Microdischarge Decay Measurement)
| P# | Hybrid | Date tested | PASS/HOLD | I-V | [7 empty] | Gold |  |  | Quality CLASS |  |  | Sum 4 sens | Sum 4 sens | Module | Module | Time to | Base | Plotted | Comment
[5 empty] | Score | [7 empty] | Stars |  |  | 2-3 | 4-5 | 6+ |  | (after MDM) | afterbonding | Post-MDM | baseline | current
P | 2 | 008 | 2004-01-18 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.481 |  | 2.699 | 1.11 | >12 hrs | 1 | yes | Ohmic IV, I > 1 ua
P | 4 | 009 | 2003-12-13 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.74 |  | 1.254 | 1.624 | NA | 1.621 | yes | Ohmic IV, I slowly rose over 12 hrs.
P | 3 | 010 | 2003-12-07 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.515 |  | 21.3 | 0.391 | 1.5 hrs | 0.37 | yes
P | 6 | 011 | 2003-12-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.414 |  | 2.317 | 0.836 | >12 hrs | 0.8 | yes | slope OK
P | 7 | 012 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.511 |  |  |  | yes
P | 49 | 013 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.447 |  | 0.42 |  |  |  | yes
P | 8 | 014 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | No data |  | 0.917 |  |  |  | yes
P | 14 | 015 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.46 |  | 0.558 |  |  |  | yes
P | 9 | 016 | 2003-12-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.372 |  | 18.608 | 0.393 | 2 hrs | 0.357 | yes
P | 10 | 017 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.332 |  | 0.672 |  |  |  | yes
P | 11 | 018 | 2003-01-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.373 |  | 0.738 |  |  |  | yes | not retested, at CERN
P | 12 | 019 | 2003-12-27 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.386 |  | 5.87 | 0.403 | 8 hrs | 0.405 | yes
P | 15 | 020 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.416 |  | 0.639 |  |  |  | yes
P | 231 | 021 | 2004-01-09 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.672 |  | 4.504 | 0.519 | 5 min | 0.442 | yes
P | 182 | 022 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.37 |  | 0.43 |  |  |  | yes
P | 13 | 023 | 2003-12-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.386 |  | 1.635 | 0.488 | 6 hrs | 0.48 | yes
P | 20 | 024 | 2003-12-04 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.331 |  | 6.66 | 0.351 | 15 min | 0.32 | yes
P | 24 | 025 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.423 |  | 0.755 |  |  |  | yes
P | 51 | 026 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.485 |  | 0.412 |  |  |  | yes
P | 32 | 027 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.368 |  | 0.352 |  |  |  | yes
P | 18 | 028 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.397 |  | 0.387 |  |  |  | yes
P | 22 | 029 | mech fail |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.353 | [5 empty] | yes
P | 28 | 030 | 2003-12-04 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.378 |  | 21.3 | 2.3 | >12 hrs | 2.5 | yes | Very noisy
P | 30 | 031 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.478 | [5 empty] | yes | one sensor x2 high
P | 29 | 032 | 2003-12-04 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.394 |  | 1.019 |  |  |  | yes
P | 38 | 033 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.606 |  | 0.442 |  |  |  | yes
P | 40 | 034 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.511 |  | 0.445 |  |  |  | yes
P | 137 | 035 | 2004-01-08 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.43 |  | 10.938 | 2.16 | 5 min | 1.43 | yes | ohmic IV
P | 42 | 036 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.412 |  | 0.373 |  |  |  | yes
P | 16 | 037 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.349 |  | 0.463 |  |  |  | yes
P | 17 | 038 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.384 |  | 0.391 |  |  |  | yes
P | 140 | 039 | 2004-01-20 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.491 |  | 6.24 | 0.454 | 1 hr | 0.365 | yes | slope 10 na/hr
P | 19 | 040 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.386 |  | 0.355 |  |  |  | yes | 8 sec wait eliminated MD
P | 23 | 041 | 2003-02-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.333 |  | 0.706 |  |  |  | yes | not retested, at CERN
P | 27 | 042 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.345 |  | 0.35 |  |  |  | yes
P | 31 | 043 | 2003-03-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.373 |  | 0.5 |  |  |  | yes | not retested, at CERN
P | 33 | 044 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.46 |  | 0.398 |  |  |  | yes
P | 35 | 045 | 2003-12-08 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.563 |  | 0.526 |  |  |  | yes
P | 193 | 046 | 2003-12-22 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.472 |  | 0.398 |  |  |  | yes
P | 338 | 047 | 2004-01-28 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.289 |  | 2.543 | 0.801 | >12 hrs | 0.8 | yes | noisy MDM curve
P | 43 | 048 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.468 |  | 0.368 |  |  |  | yes
P | 36 | 049 | 2003-12-20 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 2.86 |  | 7.53 | 0.505 | 1 hr | 0.45 | yes
P | 294 | 050 | 2004-02-13 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 8.593 | 1.798 | 11.91 | 0.611 | 30 min | 0.586 | yes | One high sensor, but does condition
P | 37 | 051 | 2003-12-25 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.588 |  | 2.97 | 0.528 | 1 hr | 0.47 | yes
P | 129 | 052 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.439 |  | 0.365 |  |  |  | yes
P | 41 | 053 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.435 |  | 0.375 |  |  |  | yes
P | 194 | 054 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.317 |  | 0.293 |  |  |  | yes
P | 45 | 055 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.496 | [5 empty] | yes
P | 143 | 056 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.364 |  | 0.474 |  |  |  | yes
P | 52 | 057 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.345 |  | 0.592 |  |  |  | yes
P | 186 | 058 | 2003-12-23 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 1.37 | 1.14 | 0.84 |  |  |  | yes
P | 167 | 059 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.405 |  | 0.4 |  |  |  | yes
P | 119 | 060 | 2004-01-06 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.604 |  | 1.209 | 1.499 | NA | 1.55 | yes | No MD, ohmic IV
P | 39 | 061 | 2003-12-27 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.494 |  | 5.034 | 0.496 | 10 min | 0.44 | yes
P | 46 | 062 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.462 |  | 0.442 |  |  |  | yes
P | 44 | 063 | 2003-12-09 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.497 |  | 2.97 | 0.474 | 12 min | 0.385 | yes
P | 331 | 064 | 2004-01-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.446 |  | 0.457 |  |  |  | yes
P | 47 | 065 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.455 |  | 0.507 |  |  |  | yes
P | 108 | 066 | 2004-01-06 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.422 |  | 2.516 | 0.342 | 30 min | 0.305 | yes | rapid drop in 6 min
P | 113 | 067 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.62 |  | 0.372 |  |  |  | yes
P | 151 | 068 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.414 |  | 0.596 |  |  |  | yes
P | 34 | 069 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.558 |  | 0.485 |  |  |  | yes
P | 48 | 070 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.435 |  | 0.395 |  |  |  | yes
P | 53 | 071 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.372 |  | 0.417 |  |  |  | yes
P | 50 | 072 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.523 |  | 0.52 |  |  |  | yes
P | 79 | 073 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.486 |  | 0.447 |  |  |  | yes
P | 55 | 074 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.404 |  | 0.482 |  |  |  | yes
P | 54 | 075 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 |  | 0.716 |  |  |  | yes
P | 56 | 076 | 2003-12-25 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.363 |  | 5.4 | 1.755 | >12 hrs | 1.87 | yes | Ohmic IV, I>1ua
P | 59 | 077 | 2003-12-22 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.356 |  | 0.65 |  |  |  | yes
P | 60 | 078 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.425 |  | 0.37 |  |  |  | yes
P | 62 | 079 | 2003-12-22 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.392 |  | 6.38 | 0.778 | 2 hrs | 0.77 | yes | bit noisy on base
P | 57 | 080 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.409 |  | 0.72 |  |  |  | yes
P | 63 | 081 | 2003-12-22 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.355 |  | 21 | 21 | >12 hrs | 21 ua | yes | 21 ua at 200 V
P | 61 | 082 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.339 |  | 0.351 |  |  |  | yes
P | 64 | 083 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.403 |  | 0.386 |  |  |  | yes
P | 68 | 084 | 2003-12-26 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.364 |  | 0.422 |  |  |  | yes
P | 72 | 085 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.513 |  | 0.449 |  |  |  | yes
P | 75 | 086 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.53 |  | 0.443 |  |  |  | yes | 8 sec wait needed
P | 98 | 087 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.517 |  | 0.427 |  |  |  | yes
P | 77 | 088 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.495 |  | 0.413 |  |  |  | yes
P | 65 | 089 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.438 |  | 0.399 |  |  |  | yes
P | 70 | 090 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.492 |  | 0.388 |  |  |  | yes
P | 71 | 091 | Mech hold |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 |  | 0.496 |  |  |  | yes
P | 73 | 092 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.457 |  | 0.491 |  |  |  | yes
P | 87 | 093 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.489 |  | 0.402 |  |  |  | yes
P | 84 | 094 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.67 |  | 0.475 |  |  |  | yes
P | 67 | 095 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.312 |  | 0.257 |  |  |  | yes | 8 sec wait killed MD
P | 107 | 096 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.455 |  | 0.345 | 0.367 | NA | 0.312 | yes | 12 hr test-normal
P | 99 | 097 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.503 |  | 0.42 |  |  |  | yes
P | 303 | 098 | 2004-01-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.577 |  | 0.595 |  |  |  | yes
P | 253 | 099 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.44 |  | 0.467 |  |  |  | yes
P | 302 | 100 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.491 |  | 0.523 |  |  |  | yes
P | 128 | 101 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.444 |  | 0.38 |  |  |  | yes
P | 139 | 102 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.417 |  | 0.782 |  |  |  | yes
P | 135 | 103 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.425 |  | 0.399 |  |  |  | yes
P | 136 | 104 | 2004-01-08 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.428 |  | 7.438 | 0.415 | 5 min | 0.379 | yes
P | 123 | 105 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.448 |  | 0.336 |  |  |  | yes
P | 157 | 106 | 2004-01-09 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.347 |  |  |  | yes | 2 sensors w abnormal IV
P | 314 | 107 | 2004-01-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.326 |  | 0.371 |  |  |  | yes
P | 175 | 108 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.385 |  | 0.399 |  |  |  | yes
P | 336 | 109 | 2004-01-27 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.325 |  | 0.387 |  |  |  | yes
P | 92 | 110 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.469 |  | 0.391 |  |  |  | yes
P | 112 | 111 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.523 |  | 0.318 |  |  |  | yes
P | 142 | 112 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.496 |  | 0.422 |  |  |  | yes | 8 sec IV killed MD
P | 120 | 113 | 2004-01-06 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 |  | 0.636 |  |  |  | yes | slight MD, pass on 8 sec
P | 111 | 114 | 2004-01-19 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.438 |  | 1.044 | 1.63 | NA | 1.67 | yes | MDM rises
P | 82 | 115 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.506 |  | 2.473 | 0.494 | >12 hrs | 0.49 | yes | fast drop 1st hr, slope OK
P | 147 | 116 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.516 |  | 0.534 |  |  |  | yes
P | 304 | 117 | 2004-01-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.533 |  | 0.567 |  |  |  | yes
P | 94 | 118 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.501 |  | 0.439 |  |  |  | yes
P | 132 | 119 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.462 |  | 0.438 |  |  |  | yes
P | 88 | 120 | 2003-12-23 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.421 |  | 0.894 |  |  |  | yes
P | 101 | 121 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.419 |  | 0.323 |  |  |  | yes
P | 156 | 122 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.383 |  | 0.338 |  |  |  | yes
P | 95 | 123 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.569 |  | 0.42 |  |  |  | yes
P | 300 | 124 | 2004-01-06 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 23.809 | 0.922 | 22.209 | 0.611 | >12 hrs | 0.55 | yes | still shows MD after 12 hrs
|  | 125 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 154 | 126 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.359 |  | 0.329 |  |  |  | yes
|  | 127 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
|  | 128 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 69 | 129 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.52 |  | 1.124 | 0.563 | >12 hrs | 0.55 | yes | slope OK
P | 144 | 130 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.417 |  | 0.349 |  |  |  | yes
P | 160 | 131 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.31 |  | 0.313 |  |  |  | yes
P | 102 | 132 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.471 |  | 0.331 |  |  |  | yes
P | 148 | 133 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.407 |  | 0.449 |  |  |  | yes
P | 145 | 134 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.386 |  | 0.485 |  |  |  | yes
P | 146 | 135 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.54 |  | 0.488 |  |  |  | yes
P | 150 | 136 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.368 |  | 0.334 |  |  |  | yes
P | 149 | 137 | 2004-01-12 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.393 |  | 0.373 |  |  |  | yes | 8 sec IV killed most MD
P | 155 | 138 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 | 1 |  |  | 0.372 |  | 0.33 |  |  |  | yes
P | 96 | 139 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.576 |  | 0.438 |  |  |  | yes
P | 124 | 140 | 2003-12-02 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.478 |  | 21.4 | 0.349 | 32 min | 0.32 | yes
P | 152 | 141 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.346 |  | 0.322 |  |  |  | yes
P | 97 | 142 | 2004-01-01 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.484 |  | 1.27 | 0.674 | >12 hrs | 0.65 | yes | slightly irregular decay curve
P | 118 | 143 | 2004-01-06 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.612 |  | 1.457 | 0.751 | >12 hrs | 0.75 | yes | slope OK
P | 83 | 144 | 2004-01-06 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.628 |  | 7.61 | 0.482 | >12 hrs | 0.48 | yes | slope OK
P | 245 | 145 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.394 |  | 0.422 |  |  |  | yes
P | 78 | 146 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.447 |  | 1.386 | 0.513 | 3.5 hrs | 0.51 | yes | T1 5min; thermistors ng
P | 76 | 147 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.769 | [5 empty] | yes | 1 sensor I(500)=.43ua
P | 326 | 148 | 2004-01-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.409 |  | 0.403 |  |  |  | yes
P | 74 | 149 | 2003-12-26 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.434 |  | 9.849 | 0.407 | 5 min | 0.355 | yes
P | 301 | 150 | 2004-01-12 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 9.957 | 0.597 | 21 | 0.526 | 4 hrs | 0.5 | yes | 21 ua at 450 V pre MDM
P | 104 | 151 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.442 |  | 0.364 |  |  |  | yes
P | 105 | 152 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.406 |  | 0.358 |  |  |  | yes
P | 90 | 153 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.428 |  | 0.429 |  |  |  | yes
P | 138 | 154 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.408 |  | 0.411 |  |  |  | yes
P | 244 | 155 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.38 |  | 0.403 |  |  |  | yes
P | 232 | 156 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.645 |  | 0.598 |  |  |  | yes
P | 127 | 157 | 2003-12-29 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.537 |  | 2.13 | 2.833 | NA | 2.8 | yes | Ohmic IV
P | 234 | 158 | 2004-01-09 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 2.712 | 1.942 | 0.93 |  |  |  | yes | 2 bad sensors
P | 80 | 159 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.468 |  | 0.389 |  |  |  | yes
P | 263 | 160 | 2003-12-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.343 |  | 0.342 |  |  |  | yes
P | 162 | 161 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.305 |  | 0.309 |  |  |  | yes
P | 125 | 162 | 2003-11-27 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.459 |  | 4.4 | 0.315 | 30 min | 0.3 | yes
P | 337 | 163 | 2004-01-28 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.497 | 0.491 | 0.513 |  |  |  | yes
P | 130 | 164 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.482 |  | 0.406 |  |  |  | yes
P | 335 | 165 | 2004-01-27 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.389 |  | 0.477 |  |  |  | yes
P | 197 | 166 | 2003-12-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.341 |  | 0.31 |  |  |  | yes
|  | 167 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 198 | 168 | 2003-12-21 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.312 |  | 1.389 | 0.302 | 15 min | 0.297 | yes
P | 163 | 169 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | inc data |  | 0.285 |  |  |  | yes | missing sensor data, OK
P | 327 | 170 | 2004-01-27 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.347 |  | 0.371 |  |  |  | yes | slight MD, w 8sec wait
P | 349 | 171 | 2004-02-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.471 |  | 0.504 |  |  |  | yes
|  | 172 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 230 | 173 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.665 |  | 0.561 |  |  |  | yes
P | 324 | 174 | 2004-01-23 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 21.624 | 0.607 | 15.75 | 0.524 | 1.2 hrs | 0.475 | yes
P | 330 | 175 | 2004-01-23 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 1.638 | 0.952 | 1.429 | 0.636 | >12 hr | 0.62 | yes | one bad sensor,slope OK
P | 249 | 176 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.463 |  | 0.453 |  |  |  | yes
P | 332 | 177 | 2004-01-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.677 | 0.634 | 0.778 |  |  |  | yes
P | 91 | 178 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 |  | 11.786 | 0.647 | >12 hrs | 0.647 | yes | big drop in 5 min, slope OK
P | 117 | 179 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.66 |  | 0.396 |  |  |  | yes
P | 177 | 180 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.344 |  | 0.433 |  |  |  | yes
|  | 181 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 134 | 182 | 2004-01-12 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.397 |  |  |  | yes | 8 sec wait killed most MD
|  | 183 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 133 | 184 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.477 |  | 0.419 |  |  |  | yes
P | 328 | 185 | 2004-01-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.38 |  | 0.433 |  |  |  | yes | Bad meas sens 4, extrap sum
P | 179 | 186 | 2003-12-23 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 8.797 | 0.411 | 6.656 | 0.38 | 1.1 hrs | 0.343 | yes
P | 153 | 187 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.38 |  | 0.386 |  |  |  | yes
P | 159 | 188 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.371 |  | 0.323 |  |  |  | yes
P | 322 | 189 | 2004-01-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 2 |  | 0.712 |  | 0.797 |  |  |  | yes
|  | 190 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 158 | 191 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.353 |  | 0.331 |  |  |  | yes
|  | 192 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 161 | 193 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.317 |  | 0.315 |  |  |  | yes
|  | 194 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 168 | 195 | Fail-accident |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.563 | 0.532 |  |  |  |  | yes | 1 abnormal sensor
P | 169 | 196 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.46 |  | 0.417 |  |  |  | yes
P | 299 | 197 | 2004-01-08 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 4.493 | 0.734 | 1.481 | 0.567 | 1 hr | 0.548 | yes | lgst drop 10 min
P | 170 | 198 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.388 |  | 0.399 |  |  |  | yes
P | 173 | 199 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 |  | 0.522 |  |  |  | yes
|  | 200 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 171 | 201 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.348 |  | 0.361 |  |  |  | yes
P | 164 | 202 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.331 |  | 0.378 |  |  |  | yes
P | 165 | 203 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.399 |  | 0.648 |  |  |  | yes | ohmic, no MD
P | 166 | 204 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.37 |  | 0.482 |  |  |  | yes
|  | 205 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 195 | 206 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.308 |  | 0.3 |  |  |  | yes
P | 172 | 207 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.435 |  | 0.357 |  |  |  | yes
P | 178 | 208 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.325 |  | 0.35 |  |  |  | yes
P | 180 | 209 | mech fail |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.367 | [5 empty] | yes
P | 204 | 210 | Mech hold |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 2.445 | 0.524 |  |  |  |  | yes | one sensor, high, seems to condition OK
P | 185 | 211 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.448 | 0.393 | 0.401 |  |  |  | yes
P | 323 | 212 | 2004-01-22 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 2 |  | 0.68 |  | 0.699 |  |  |  | yes | slight MD, w 8sec wait
P | 214 | 213 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.63 |  | 0.465 |  |  |  | yes
P | 191 | 214 | 2003-12-23 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.485 |  | 2.639 | 0.383 | 30 min | 0.362 | yes
P | 333 | 215 | 2004-01-27 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.47 |  | 0.525 |  |  |  | yes
P | 200 | 216 | PA problem |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.343 | [5 empty] | yes
P | 215 | 217 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.642 |  | 0.497 |  |  |  | yes
P | 199 | 218 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.377 |  | 0.32 |  |  |  | yes
P | 202 | 219 | 2003-12-23 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.388 |  | 3.439 | 0.358 | 45 min | 0.338 | yes
P | 203 | 220 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 | 1 |  |  | 0.459 |  | 0.348 |  |  |  | yes
P | 288 | 221 | 2004-01-07 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.346 |  | 0.353 |  |  |  | yes
P | 219 | 222 | 2004-01-09 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 13.993 | 0.399 | 10.777 | 0.378 | 8 min | 0.34 | yes
P | 209 | 223 | 2004-01-12 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.572 |  | 10.193 | 0.446 | 25 min | 0.382 | yes
P | 208 | 224 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.525 |  | 0.392 |  |  |  | yes
|  | 225 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 181 | 226 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.369 |  | 0.352 |  |  |  | yes
P | 196 | 227 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.37 |  | 0.312 |  |  |  | yes
|  | 228 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 183 | 229 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.391 |  | 0.399 |  |  |  | yes
P | 187 | 230 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.398 |  | 0.316 |  |  |  | yes
P | 375 | 231 | 2004-02-20 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.347 |  | 0.328 |  |  |  | yes
P | 184 | 232 | chipped TL |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 | [5 empty] | yes
P | 240 | 233 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.379 |  | 0.356 |  |  |  | yes
P | 216 | 234 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.642 |  | 0.458 |  |  |  | yes
P | 205 | 235 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.431 |  | 0.354 |  |  |  | yes
P | 217 | 236 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.496 |  | 0.437 |  |  |  | yes
P | 206 | 237 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.522 |  | 0.386 |  |  |  | yes
P | 250 | 238 | 2003-11-26 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 4 |  |  | 9.5 | 1.85 | 0.628 |  |  |  | yes | 1.8 ua on 1 s wait
P | 220 | 239 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.544 |  | 0.451 |  |  |  | yes
P | 207 | 240 | 2004-01-12 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.526 |  | 7.33 | 0.406 | 8 min | 0.36 | yes
P | 225 | 241 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.504 |  | 0.388 |  |  |  | yes
P | 221 | 242 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.393 |  | 0.357 |  |  |  | yes
P | 251 | 243 | 2003-11-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.462 |  | 0.475 |  |  |  | yes
P | 218 | 244 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.446 |  | 0.474 |  |  |  | yes
P | 248 | 245 | 2003-11-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.456 |  | 0.482 |  |  |  | yes
P | 329 | 246 | 2004-01-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.414 |  | 0.466 |  |  |  | yes
P | 318 | 247 | 2004-01-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 0.648 | 0.528 | 0.565 |  |  |  | yes
P | 305 | 248 | 2004-01-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 0.569 |  | 0.614 |  |  |  | yes
P | 227 | 249 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.552 |  | 0.414 |  |  |  | yes
P | 211 | 250 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.469 |  | 0.389 |  |  |  | yes
P | 367 | 251 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.363 |  | 0.42 |  |  |  | yes
P | 264 | 252 | 2003-12-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.347 |  | 0.35 |  |  |  | yes
P | 266 | 253 | 2003-12-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.32 |  | 0.336 |  |  |  | yes
P | 226 | 254 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.528 |  | 0.395 |  |  |  | yes
P | 213 | 255 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.492 |  | 0.447 |  |  |  | yes
P | 246 | 256 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.465 |  | 0.484 |  |  |  | yes
P | 247 | 257 | 2003-11-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.451 |  | 0.458 |  |  |  | yes
P | 210 | 258 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.491 |  | 0.393 |  |  |  | yes
P | 235 | 259 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.475 |  | 0.416 |  |  |  | yes
P | 262 | 260 | 2003-12-10 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.318 |  | 13.81 | 0.33 | 5 min | 0.28 | yes
|  | 261 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 212 | 262 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.409 |  | 0.509 |  |  |  | yes
P | 228 | 263 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.51 |  | 0.415 |  |  |  | yes
P | 343 | 264 | 2004-01-30 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.339 |  | 0.371 |  |  |  | yes
P | 242 | 265 | 2004-01-12 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 21.505 | 0.401 | 5.891 | 0.421 | 40 min | 0.35 | yes
P | 239 | 266 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.42 |  | 0.455 |  |  |  | yes
P | 236 | 267 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.489 |  | 0.424 |  |  |  | yes
P | 241 | 268 | 2003-12-01 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.467 | 0.385 | 21.4 | 4.99 | >12 hrs | 4.8 | yes | "base" 4.8 ua in 2 hrs
P | 313 | 269 | 2001-01-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.43 |  | 0.422 |  |  |  | yes | one sensor w unusual IV
P | 238 | 270 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.398 |  | 0.394 |  |  |  | yes | 8sec IV killed MD
P | 269 | 271 | 2003-12-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.506 |  | 0.479 |  |  |  | yes
P | 268 | 272 | 2003-12-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.518 |  | 0.473 |  |  |  | yes
P | 283 | 273 | 2004-01-01 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.54 |  | 1.366 | 1.769 | 10 min | 1.46 | yes | erratic, MDM start 21 ua
P | 229 | 274 | 2004-01-10 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.492 |  | 21 | 0.414 | 3 hrs | 0.395 | yes | good conditioning curve
P | 243 | 275 | 2004-01-06 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 1.156 | 0.382 | 5.872 | 0.459 | 8 min | 0.42 | yes
P | 312 | 276 | 2004-01-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 0.545 |  | 0.568 |  |  |  | yes
P | 267 | 277 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.485 |  | 0.537 |  |  |  | yes
|  | 278 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 279 | 279 | 2003-12-18 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 6.703 | 0.443 | 12.063 | 0.394 | 20 min | 0.342 | yes
P | 272 | 280 | 2003-12-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.546 |  | 0.499 |  |  |  | yes
P | 317 | 281 | 2004-01-16 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 15.301 | 5.409 | 21.143 | 0.403 | 3 hrs | 0.38 | yes
|  | 282 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 387 | 283 | 2004-02-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 2 |  |  | 4.809 | 0.703 | 11.915 | 0.679 | 1.8 hrs | 0.625 | yes | Long decay from bonding, not sensors
P | 376 | 284 | 2004-02-23 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.34 |  | 0.866 |  |  |  | yes
P | 278 | 285 | 2003-12-20 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.391 |  | 4.418 | 0.491 | 2 hrs | 0.49 | yes | Bad initial IV ??
|  | 286 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 252 | 287 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.447 |  | 0.477 |  |  |  | yes
P | 256 | 288 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.613 |  | 0.558 |  |  |  | yes
P | 255 | 289 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.589 |  | 0.586 |  |  |  | yes | early IV 21 ua sens 4
P | 259 | 290 | 2004-01-12 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.649 |  | 1.219 | 1.373 | NA | 1.37 | yes | normal IV, just high
P | 310 | 291 | 2004-01-14 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.445 |  | 0.573 |  |  |  | yes
P | 260 | 292 | 2004-01-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.565 |  | 0.592 |  |  |  | yes
P | 341 | 293 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.358 |  | 0.373 |  |  |  | yes
P | 306 | 294 | 2004-01-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.521 |  | 0.489 |  |  |  | yes
P | 309 | 295 | 2004-01-14 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.461 |  | 0.516 |  |  |  | yes
P | 307 | 296 | 2004-01-15 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.425 |  | 0.571 |  |  |  | yes
P | 354 | 297 | 2004-02-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.4 |  | 0.432 |  |  |  | yes
P | 261 | 298 | 2003-12-08 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.485 |  | 0.579 |  |  |  | yes | 0.1 ua MD
P | 257 | 299 | 2003-12-04 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 10.5 | 0.736 | 2.79 | 0.999 | >12 hrs | 1 | yes | slope OK
P | 321 | 300 | 2004-01-22 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 0.516 |  | 3.574 | 0.49 | 2.5 hrs | 0.465 | yes
P | 281 | 301 | 2003-12-19 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.431 |  | 0.865 |  |  |  | yes
P | 355 | 302 | 2004-02-10 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 4.255 | 3.293 | 1.701 | 1.325 | 12 hrs | 1.2 | yes | Acceptable slope in 16 hrs
P | 265 | 303 | 2003-12-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.339 |  | 0.37 |  |  |  | yes
P | 282 | 304 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.375 |  | 0.398 |  |  |  | yes
P | 258 | 305 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.612 |  | 0.64 |  |  |  | yes | sensor lkge high
P | 274 | 306 | 2003-12-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.482 |  | 0.398 |  |  |  | yes
P | 273 | 307 | 2003-12-16 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.504 |  | 11.492 | 0.452 | 5 min | 0.375 | yes
P | 308 | 308 | 2004-01-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.427 |  | 0.472 |  |  |  | yes
P | 315 | 309 | 2004-01-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.406 |  | 0.465 |  |  |  | yes
P | 270 | 310 | 2003-12-15 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.504 |  | 0.487 |  |  |  | yes
P | 277 | 311 | 2003-12-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.547 |  | 0.51 |  |  |  | yes
P | 311 | 312 | 2004-01-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.42 |  | 0.453 |  |  |  | yes
P | 276 | 313 | 2003-12-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.506 |  | 0.452 |  |  |  | yes
P | 342 | 314 | 2004-01-30 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.348 |  | 0.353 |  |  |  | yes
P | 280 | 315 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.448 |  | 0.44 |  |  |  | yes
|  | 316 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 284 | 317 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.9 |  | 0.921 |  |  |  | yes
|  | 318 |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 285 | 319 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.965 |  | 0.925 |  |  |  | yes
P | 287 | 320 | 2004-01-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.392 |  | 1.936 | 0.786 | >12 hrs | 0.78 | yes | slope OK
P | 275 | 321 | 2003-12-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.487 |  | 0.539 |  |  |  | yes
P | 286 | 322 | 2004-01-02 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.867 |  |  |  | yes | ohmic IV curve
P | 289 | 323 | 2004-01-07 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.802 |  | 1.461 | 0.966 | 3 hrs | 0.95 | yes
P | 290 | 324 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.872 |  | 0.957 |  |  |  | yes
P | 292 | 325 | 2004-01-15 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.764 |  | 0.977 |  |  |  | yes
P | 296 | 326 | 2004-01-16 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 3 |  | 30.9 | 3.533 | 8.814 | 0.472 | 4 hrs | 0.46 | yes | rapid drop in 1st 6 min
P | 291 | 327 | 2004-01-12 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.789 |  | 1.865 | 1.176 | 1.5 hrs | 1.116 | yes | OK, just high
P | 340 | 328 | 2004-01-28 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.377 |  | 0.361 |  |  |  | yes
P | 334 | 329 | 2004-01-28 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.459 |  | 0.47 |  |  |  | yes
P | 386 | 330 | 2004-02-25 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 4 |  |  | 0.52 |  | 0.625 |  |  |  | yes
P | 371 | 331 | 2004-02-19 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.362 |  | 0.394 |  |  |  | yes
P | 370 | 332 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.459 |  | 0.512 |  |  |  | yes
P | 353 | 333 | 2004-02-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.398 |  | 0.433 |  |  |  | yes
P | 352 | 334 | 2004-02-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.369 |  | 0.391 |  |  |  | yes
P | 377 | 335 | 2004-02-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.298 |  | 0.309 |  |  |  | yes
P | 365 | 336 | 2004-02-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.428 |  | 0.465 |  |  |  | yes
P | 363 | 337 | 2004-02-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.381 |  | 0.417 |  |  |  | yes
P | 379 | 338 | 2004-02-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.333 |  | 0.345 |  |  |  | yes
P | 396 | 339 | 2004-03-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.457 |  | 0.468 |  |  |  | yes
P | 362 | 340 | 2004-02-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.35 |  | 0.412 |  |  |  | yes
341
P | 369 | 342 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.344 |  | 0.396 |  |  |  | yes
343
344
345
346
347
348
P | 351 | 349 | 2004-02-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.382 |  | 0.384 |  |  |  | yes
350
351
352
P | 359 | 353 | 2004-02-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.348 |  | 0.385 |  |  |  | yes
P | 358 | 354 | 2004-02-10 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.59 | 0.398 | 0.553 |  |  |  | yes | 8 sec wait
P | 356 | 355 | 2004-02-10 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 2 |  | 18.635 | 3.027 | 6.739 | 0.349 | 2.5 hrs | 0.349 | yes | very well behaved
P | 384 | 356 | 2004-02-25 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 11.243 | 0.607 | 0.766 |  |  |  | yes
P | 350 | 357 | 2004-02-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.376 |  | 0.411 |  |  |  | yes
358
P | 391 | 359 | 2004-02-27 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.411 |  | 21.4 | 21.1 | >12 hrs | 21 | yes | 21 ua at 200V, at 450 V after 12 hrs
360
P | 361 | 361 | 2004-02-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.358 |  | 0.385 |  |  |  | yes
P | 381 | 362 | 2004-02-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.468 |  | 0.465 |  |  |  | yes
P | 373 | 363 | 2004-02-19 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  | 1 |  | 0.4 |  | 0.403 |  |  |  | yes
364
P | 348 | 365 | 2004-02-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.382 |  | 0.37 |  |  |  | yes
P | 368 | 366 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.367 |  | 0.418 |  |  |  | yes
367
368
369
370
P | 385 | 371 | 2004-02-25 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 18.498 | 1.051 | 11.481 | 0.663 | 1.15 hrs | 0.63 | yes | compound MDM slope from different sensors
372
P | 383 | 373 | 2004-02-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.485 |  | 0.51 |  |  |  | yes
P | 392 | 374 | 2004-03-03 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.482 |  | 0.602 |  |  |  | yes | 1 abnormal sensor, afterbonding IV slightly irregular
P | 344 | 375 | 2004-02-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.284 |  | 0.322 |  |  |  | yes
P | 390 | 376 | 2004-02-27 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 5.237 | 0.424 | 11.566 | 0.373 | 30 min | 0.341 | yes | Conditions very well
P | 345 | 377 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.31 |  | 0.343 |  |  |  | yes
P | 366 | 378 | 2004-02-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.92 | 0.368 | 0.475 |  |  |  | yes | possibly erroneous hi meas
P | 346 | 379 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.328 |  | 0.354 |  |  |  | yes
P | 382 | 380 | 2004-02-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.388 |  | 0.388 |  |  |  | yes
P | 364 | 381 | 2004-02-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.359 |  | 0.374 |  |  |  | yes
P | 360 | 382 | 2004-02-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.321 |  | 0.358 |  |  |  | yes
P | 347 | 383 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.368 |  | 0.411 |  |  |  | yes
P | 374 | 384 | 2004-02-20 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.402 |  | 0.394 |  |  |  | yes
P | 372 | 385 | 2004-02-19 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.354 |  | 0.382 |  |  |  | yes
386
387
388
389
390
391
392
393
394
P | 395 | 395 | 2004-03-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.596 | 0.481 | 0.608 |  |  |  | yes
P | 389 | 396 | 2004-02-28 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 16.825 | 0.487 | 21.22 | 0.451 | 25 min | 0.41 | yes | Conditions very well
P | 325 | 397 | 2004-03-05 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 3.118 | 1.184 | 1.3 | 0.491 | 8 hrs | 0.51 | yes | One bad sensor, seems to respond to conditioning
398
P | 399 | 399 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.438 |  | 0.486 |  |  |  | yes
P | 400 | 400 | 2004-03-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.438 |  | 0.472 |  |  |  | yes
P | 401 | 401 | 2004-03-08 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.369 |  | 1.814 | 0.963 | 12 hrs | 0.963 | yes | ohmic IV curve, noisy MDM
P | 397 | 402 | 2004-03-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.47 |  | 0.506 |  |  |  | yes
P | 393 | 403 | 2004-03-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.666 | 3.14 | 3.485 | 0.62 | 12 hrs | 0.63 | yes | OK to bond
P | 319 | 404 | 2004-03-05 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 10.035 | 5.384 | 5.791 | 0.501 | 24 hrs | 0.5 | yes | Bases in 24 hours
P | 398 | 405 | 2004-03-05 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 13.947 | 0.599 | 6.763 | 0.456 | 15 min | 0.404 | yes | Very fine conditioning
406
407
P | 402 | 408 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.374 |  | 0.427 |  |  |  | yes
409
410
411
412
P | 406 | 413 | 2004-03-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.427 |  | 0.489 |  |  |  | yes
P | 404 | 414 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.468 |  | 0.539 |  |  |  | yes
P | 405 | 415 | 2004-03-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.434 |  | 0.472 |  |  |  | yes
P | 403 | 416 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.392 |  | 0.379 |  |  |  | yes
P | 408 | 417 | 2004-03-11 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 9.325 | 0.482 | 15.858 | 0.459 | 45 min | 0.42 | yes | noisy MDM curve
P | 409 | 418 | 2004-03-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 |  | 0.532 |  |  |  | yes
P | 407 | 419 | 2004-03-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.449 |  | 0.496 |  |  |  | yes
P | 411 | 420 | 2004-03-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.465 |  | 0.549 |  |  |  | yes
421
422
423
424
425
426
P | 410 | 427 | 2004-03-11 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.581 | 6.631 | 12.05 | 0.504 | 25 min | 0.502 | yes | break @ 300, but does condition
428
429
430
P | 122 | 980 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.463 |  | 0.332 |  |  |  | yes | 8 sec
P | 114 | 981 | Rework |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.612 | [5 empty] | yes
P | 115 | 983 | 2004-01-06 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.626 |  | 0.738 |  |  |  | yes | IV ~linear, but <1ua
P | 121 | 993 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.446 |  | 0.337 |  |  |  | yes
P | 1 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.508 | [5 empty] | yes
P | 5 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.56 | [5 empty] | yes
P | 21 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.332 | [5 empty] | yes
P | 25 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 98.3 | [5 empty] | yes | One really bad sensor
P | 26 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.344 | [5 empty] | yes
P | 58 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.422 | [5 empty] | yes | one sensor data missing, entry from logbook
P | 66 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.429 | [5 empty] | yes | 1 sensor slight MD
P | 81 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.546 | [5 empty] | yes
P | 85 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.451 | [5 empty] | yes
P | 86 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.401 | [5 empty] | yes
P | 89 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 | [5 empty] | yes
P | 93 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.465 | [5 empty] | yes
P | 100 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.639 | [5 empty] | yes | 1 sensor I(500) = .30 ua
P | 103 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.389 | [5 empty] | yes
P | 106 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.467 | [5 empty] | yes
P | 109 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.468 | [5 empty] | yes
P | 110 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.488 | [5 empty] | yes
P | 116 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.716 | [5 empty] | yes | 2 abnormal sensors
P | 126 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | >300 | [5 empty] | yes
P | 131 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | No IV data | [5 empty] | yes
P | 141 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.384 | [5 empty] | yes | afterbonding.iv file belongs to P144
P | 174 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 5.036 | 0.931 |  |  |  |  | yes | One bad sensor, seems to condition
P | 176 |  | RELEASE |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.759 | 0.403 | [5 empty] | One bad sensor, seems to condition, LOGBOOK 9/15/03 indicates conditioning problems
P | 188 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.355 | [5 empty] | yes
P | 189 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.522 | [5 empty] | yes
P | 190 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.488 | [5 empty] | yes
P | 192 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 385 | [5 empty] | yes | One insanely high sensor
P | 201 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 73.7 | 0.346 |  |  |  |  | yes | One bad sensor, seems to condition
P | 222 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.492 | [5 empty] | yes
P | 223 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 347 | 0.445 |  |  |  |  | yes | One chipped sensor, remeas next day was OK??
P | 224 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.539 | 0.474 |  |  |  |  | yes
P | 233 |  | No hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.659 | [5 empty] | yes
P | 237 |  | shipped |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | inc data | [5 empty] | yes | "BURNED"
P | 254 |  | no hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.489 | [5 empty] | yes
P | 271 |  | no hybrid |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.523 | [5 empty] | yes
P | 293 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 23.238 | 17.531 |  |  |  |  | yes | 3 abnormal sensors, Nobu suggested also to hold
P | 295 |  | RELEASE |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 3 |  | 30.33 | 22.07 | [5 empty] | One bad sensor, seems to improve with conditioning
P | 297 |  | release |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 32.66 | 18.46 |  |  |  |  | yes | One bad sensor, seems to improve with conditioning
P | 298 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 3 |  | 39.453 | 21.975 |  |  |  |  | yes | One very bad sensor, Nobu suggested holding
P | 316 |  | Hold |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 1.331 | 1.065 |  |  |  |  | yes | 2 high sensors
P | 320 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 2 |  | 2.782 | 2.77 |  |  |  |  | yes | One bad sensor, no conditioning effect
P | 339 |  | Hold |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.339 | [5 empty] | yes
P | 357 |  | RELEASE |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 2 | 1 | 17.504 | 6.127 | [5 empty] | One sensor >6ua after conditioning
P | 378 |  | HOLD |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.376 | 0.343 |  |  |  |  | yes
P | 380 |  | HOLD |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.357 | [5 empty] | yes
P | 388 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 2 | 1 | 48.378 | 26.413 |  |  |  |  | yes | May possibly condition
P | 394 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.566 | 0.493 |  |  |  |  | yes | Very noisy, NOT probe-related
P | 412 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.511
P | 413 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.514
P | 414 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.668 | 0.599 | [5 empty] | one sensor somewhat erratic
P | 415 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.404
P | 416 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.384
P | 417 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.423
P | 418 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.417
P | 419 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.476
P | 420 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.465
P | 421 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.422
P | 422 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.431
P | 423 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.454
P | 424 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.476
P | 425 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.572
P | 426 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.469
P | 427 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.399
P | 428 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.384
P | 429 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.419
P | 430 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.463
P | 431 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.493
P | 432 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 13.184 | 0.133 | [5 empty] | 1 sensor conditions
P | 433 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.403
P | 434 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.378
P | 435 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.412
P | 436 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.446
P | 437 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.463
P | 438 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.451
P | 439 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.415
P | 440 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.421
P | 441 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.44
P | 442 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.505 | [6 empty] | 1 abnormal sensor
P | 443 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.832 | 0.762 | [5 empty] | 1 abnormal sensor
P | 444 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.426
P | 445 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.436 | [6 empty] | Backside sensors both noisy
P | 446 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.512
P | 447 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.536 | [6 empty] | 1 abnormal sensor
P | 448 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.461
P | 449 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 2.477 | 0.515 | [5 empty] | 2 abnormal sensors, condition OK
P | 450 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.996 | 2.035 | [5 empty] | 2 abnormal sensors, probably will condition
P | 451 |  | IV & glue | hold iv |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.854 | 2.91 | [5 empty] | One very abnormal sensor
P | 452 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 11.859 | 0.54 | [5 empty] | 1 sens, conditions in 1 min
P | 453 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.776 | 0.599 | [5 empty] | 1 abnormal sensor, conditions ~OK
P | 454 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.508
P | 455 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.503
P | 456 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.463
P | 457 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.464
P | 458 |  |  | HOLD |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.96 | 21.91 | [5 empty] | 1 sensor, 21ua @ 30V no condition effect
P | 459 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.377
P | 460 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.467
P | 461 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.433
P | 462 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.428
P | 463 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 11.905 | 0.416 | [5 empty] | 1 sensor conditions 2.5 min
P | 464 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 465 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 466 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 467 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 468 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 469 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 470 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 471 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 472 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 473 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 474 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 475 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 476 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 477 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 478 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 479 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 480 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 481 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 482 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 483 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 484 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 485 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 486 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 487 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 488 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 489 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 490 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 491 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 492 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 493 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 494 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 495 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 496 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 497 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 498 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 499 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 500 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 501 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 502 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 503 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 504 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 505 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 506 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 507 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 508 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 509 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 510 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 511 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 512 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 513 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 514 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 515 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 516 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 517 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 518 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 519 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 520 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 521 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 522 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 523 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 524 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 525 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 526 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 527 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 528 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 529 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 530 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 531 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 532 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 533 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 534 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 535 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 536 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 537 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 538 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 539 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 540 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 541 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 542 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 543 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 544 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 545 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 546 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 547 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 548 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 549 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0
P | 550 |  |  |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0

## Sorting of problem modules
| Module ID | [18 empty] | I(500) |  |  | MDM (Microdischarge Decay Measurement)
| P# | Hybrid | Date tested | PASS/HOLD | I-V | [7 empty] | Gold |  |  | Quality CLASS |  |  | Sum 4 sens | Sum 4 sens | Module | Module | Time to | Base | Plotted | Comment
[5 empty] | Score | [7 empty] | Stars |  |  | 2-3 | 4-5 | 6+ |  | (after MDM) | afterbonding | Post-MDM | baseline | current
P | 391 | 359 | 2004-02-27 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.411 |  | 21.4 | 21.1 | >12 hrs | 21 | yes | 21 ua at 200V, at 450 V after 12 hrs
P | 63 | 081 | 2003-12-22 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.355 |  | 21 | 21 | >12 hrs | 21 ua | yes | 21 ua at 200 V
P | 241 | 268 | 2003-12-01 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.467 | 0.385 | 21.4 | 4.99 | >12 hrs | 4.8 | yes | "base" 4.8 ua in 2 hrs
P | 28 | 030 | 2003-12-04 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.378 |  | 21.3 | 2.3 | >12 hrs | 2.5 | yes | Very noisy
P | 127 | 157 | 2003-12-29 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.537 |  | 2.13 | 2.833 | NA | 2.8 | yes | Ohmic IV
P | 137 | 035 | 2004-01-08 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.43 |  | 10.938 | 2.16 | 5 min | 1.43 | yes | ohmic IV
P | 283 | 273 | 2004-01-01 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.54 |  | 1.366 | 1.769 | 10 min | 1.46 | yes | erratic, MDM start 21 ua
P | 56 | 076 | 2003-12-25 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.363 |  | 5.4 | 1.755 | >12 hrs | 1.87 | yes | Ohmic IV, I>1ua
P | 111 | 114 | 2004-01-19 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.438 |  | 1.044 | 1.63 | NA | 1.67 | yes | MDM rises
P | 4 | 009 | 2003-12-13 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.74 |  | 1.254 | 1.624 | NA | 1.621 | yes | Ohmic IV, I slowly rose over 12 hrs.
P | 119 | 060 | 2004-01-06 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.604 |  | 1.209 | 1.499 | NA | 1.55 | yes | No MD, ohmic IV
P | 259 | 290 | 2004-01-12 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.649 |  | 1.219 | 1.373 | NA | 1.37 | yes | normal IV, just high
P | 355 | 302 | 2004-02-10 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 4.255 | 3.293 | 1.701 | 1.325 | 12 hrs | 1.2 | yes | Acceptable slope in 16 hrs
P | 291 | 327 | 2004-01-12 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.789 |  | 1.865 | 1.176 | 1.5 hrs | 1.116 | yes | OK, just high
P | 2 | 008 | 2004-01-18 | PASS | 4 | 0 | 0 | 0 | 0 | 1 | 0 | 0 |  | 0 | 0 |  |  |  | 0.481 |  | 2.699 | 1.11 | >12 hrs | 1 | yes | Ohmic IV, I > 1 ua
P | 257 | 299 | 2003-12-04 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 10.5 | 0.736 | 2.79 | 0.999 | >12 hrs | 1 | yes | slope OK
P | 289 | 323 | 2004-01-07 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.802 |  | 1.461 | 0.966 | 3 hrs | 0.95 | yes
P | 401 | 401 | 2004-03-08 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.369 |  | 1.814 | 0.963 | 12 hrs | 0.963 | yes | ohmic IV curve, noisy MDM
P | 6 | 011 | 2003-12-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.414 |  | 2.317 | 0.836 | >12 hrs | 0.8 | yes | slope OK
P | 338 | 047 | 2004-01-28 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.289 |  | 2.543 | 0.801 | >12 hrs | 0.8 | yes | noisy MDM curve
P | 287 | 320 | 2004-01-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.392 |  | 1.936 | 0.786 | >12 hrs | 0.78 | yes | slope OK
P | 62 | 079 | 2003-12-22 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.392 |  | 6.38 | 0.778 | 2 hrs | 0.77 | yes | bit noisy on base
P | 118 | 143 | 2004-01-06 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.612 |  | 1.457 | 0.751 | >12 hrs | 0.75 | yes | slope OK
P | 387 | 283 | 2004-02-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 2 |  |  | 4.809 | 0.703 | 11.915 | 0.679 | 1.8 hrs | 0.625 | yes | Long decay from bonding, not sensors
P | 97 | 142 | 2004-01-01 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.484 |  | 1.27 | 0.674 | >12 hrs | 0.65 | yes | slightly irregular decay curve
P | 385 | 371 | 2004-02-25 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 18.498 | 1.051 | 11.481 | 0.663 | 1.15 hrs | 0.63 | yes | compound MDM slope from different sensors
P | 91 | 178 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 |  | 11.786 | 0.647 | >12 hrs | 0.647 | yes | big drop in 5 min, slope OK
P | 330 | 175 | 2004-01-23 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 1.638 | 0.952 | 1.429 | 0.636 | >12 hr | 0.62 | yes | one bad sensor,slope OK
P | 393 | 403 | 2004-03-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.666 | 3.14 | 3.485 | 0.62 | 12 hrs | 0.63 | yes | OK to bond
P | 300 | 124 | 2004-01-06 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 23.809 | 0.922 | 22.209 | 0.611 | >12 hrs | 0.55 | yes | still shows MD after 12 hrs
P | 69 | 129 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.52 |  | 1.124 | 0.563 | >12 hrs | 0.55 | yes | slope OK
P | 301 | 150 | 2004-01-12 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 9.957 | 0.597 | 21 | 0.526 | 4 hrs | 0.5 | yes | 21 ua at 450 V pre MDM
P | 324 | 174 | 2004-01-23 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 21.624 | 0.607 | 15.75 | 0.524 | 1.2 hrs | 0.475 | yes
P | 78 | 146 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.447 |  | 1.386 | 0.513 | 3.5 hrs | 0.51 | yes | T1 5min; thermistors ng
P | 319 | 404 | 2004-03-05 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 10.035 | 5.384 | 5.791 | 0.501 | 24 hrs | 0.5 | yes | Bases in 24 hours
P | 82 | 115 | 2003-12-26 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.506 |  | 2.473 | 0.494 | >12 hrs | 0.49 | yes | fast drop 1st hr, slope OK
P | 278 | 285 | 2003-12-20 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.391 |  | 4.418 | 0.491 | 2 hrs | 0.49 | yes | Bad initial IV ??
P | 325 | 397 | 2004-03-05 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 3.118 | 1.184 | 1.3 | 0.491 | 8 hrs | 0.51 | yes | One bad sensor, seems to respond to conditioning
P | 321 | 300 | 2004-01-22 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 0.516 |  | 3.574 | 0.49 | 2.5 hrs | 0.465 | yes
P | 13 | 023 | 2003-12-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.386 |  | 1.635 | 0.488 | 6 hrs | 0.48 | yes
P | 83 | 144 | 2004-01-06 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.628 |  | 7.61 | 0.482 | >12 hrs | 0.48 | yes | slope OK
P | 296 | 326 | 2004-01-16 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 3 |  | 30.9 | 3.533 | 8.814 | 0.472 | 4 hrs | 0.46 | yes | rapid drop in 1st 6 min
P | 229 | 274 | 2004-01-10 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.492 |  | 21 | 0.414 | 3 hrs | 0.395 | yes | good conditioning curve
P | 317 | 281 | 2004-01-16 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 15.301 | 5.409 | 21.143 | 0.403 | 3 hrs | 0.38 | yes
P | 12 | 019 | 2003-12-27 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.386 |  | 5.87 | 0.403 | 8 hrs | 0.405 | yes
P | 9 | 016 | 2003-12-03 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.372 |  | 18.608 | 0.393 | 2 hrs | 0.357 | yes
P | 3 | 010 | 2003-12-07 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.515 |  | 21.3 | 0.391 | 1.5 hrs | 0.37 | yes
P | 179 | 186 | 2003-12-23 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 8.797 | 0.411 | 6.656 | 0.38 | 1.1 hrs | 0.343 | yes
P | 356 | 355 | 2004-02-10 | PASS | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |  | 0 | 0 | 1 | 2 |  | 18.635 | 3.027 | 6.739 | 0.349 | 2.5 hrs | 0.349 | yes | very well behaved
P | 294 | 050 | 2004-02-13 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 8.593 | 1.798 | 11.91 | 0.611 | 30 min | 0.586 | yes | One high sensor, but does condition
P | 299 | 197 | 2004-01-08 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 4.493 | 0.734 | 1.481 | 0.567 | 1 hr | 0.548 | yes | lgst drop 10 min
P | 37 | 051 | 2003-12-25 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.588 |  | 2.97 | 0.528 | 1 hr | 0.47 | yes
P | 231 | 021 | 2004-01-09 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.672 |  | 4.504 | 0.519 | 5 min | 0.442 | yes
P | 36 | 049 | 2003-12-20 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 2.86 |  | 7.53 | 0.505 | 1 hr | 0.45 | yes
P | 410 | 427 | 2004-03-11 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.581 | 6.631 | 12.05 | 0.504 | 25 min | 0.502 | yes | break @ 300, but does condition
P | 39 | 061 | 2003-12-27 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.494 |  | 5.034 | 0.496 | 10 min | 0.44 | yes
P | 44 | 063 | 2003-12-09 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.497 |  | 2.97 | 0.474 | 12 min | 0.385 | yes
P | 408 | 417 | 2004-03-11 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 9.325 | 0.482 | 15.858 | 0.459 | 45 min | 0.42 | yes | noisy MDM curve
P | 243 | 275 | 2004-01-06 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 1.156 | 0.382 | 5.872 | 0.459 | 8 min | 0.42 | yes
P | 398 | 405 | 2004-03-05 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 13.947 | 0.599 | 6.763 | 0.456 | 15 min | 0.404 | yes | Very fine conditioning
P | 140 | 039 | 2004-01-20 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.491 |  | 6.24 | 0.454 | 1 hr | 0.365 | yes | slope 10 na/hr
P | 273 | 307 | 2003-12-16 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.504 |  | 11.492 | 0.452 | 5 min | 0.375 | yes
P | 389 | 396 | 2004-02-28 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 16.825 | 0.487 | 21.22 | 0.451 | 25 min | 0.41 | yes | Conditions very well
P | 209 | 223 | 2004-01-12 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.572 |  | 10.193 | 0.446 | 25 min | 0.382 | yes
P | 242 | 265 | 2004-01-12 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 21.505 | 0.401 | 5.891 | 0.421 | 40 min | 0.35 | yes
P | 136 | 104 | 2004-01-08 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.428 |  | 7.438 | 0.415 | 5 min | 0.379 | yes
P | 74 | 149 | 2003-12-26 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.434 |  | 9.849 | 0.407 | 5 min | 0.355 | yes
P | 207 | 240 | 2004-01-12 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.526 |  | 7.33 | 0.406 | 8 min | 0.36 | yes
P | 279 | 279 | 2003-12-18 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 6.703 | 0.443 | 12.063 | 0.394 | 20 min | 0.342 | yes
P | 191 | 214 | 2003-12-23 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.485 |  | 2.639 | 0.383 | 30 min | 0.362 | yes
P | 219 | 222 | 2004-01-09 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 13.993 | 0.399 | 10.777 | 0.378 | 8 min | 0.34 | yes
P | 390 | 376 | 2004-02-27 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 5.237 | 0.424 | 11.566 | 0.373 | 30 min | 0.341 | yes | Conditions very well
P | 202 | 219 | 2003-12-23 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.388 |  | 3.439 | 0.358 | 45 min | 0.338 | yes
P | 20 | 024 | 2003-12-04 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.331 |  | 6.66 | 0.351 | 15 min | 0.32 | yes
P | 124 | 140 | 2003-12-02 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.478 |  | 21.4 | 0.349 | 32 min | 0.32 | yes
P | 108 | 066 | 2004-01-06 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.422 |  | 2.516 | 0.342 | 30 min | 0.305 | yes | rapid drop in 6 min
P | 262 | 260 | 2003-12-10 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.318 |  | 13.81 | 0.33 | 5 min | 0.28 | yes
P | 125 | 162 | 2003-11-27 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.459 |  | 4.4 | 0.315 | 30 min | 0.3 | yes
P | 198 | 168 | 2003-12-21 | PASS | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.312 |  | 1.389 | 0.302 | 15 min | 0.297 | yes
P | 29 | 032 | 2003-12-04 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.394 |  | 1.019 |  |  |  | yes
P | 234 | 158 | 2004-01-09 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 2.712 | 1.942 | 0.93 |  |  |  | yes | 2 bad sensors
P | 88 | 120 | 2003-12-23 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.421 |  | 0.894 |  |  |  | yes
P | 286 | 322 | 2004-01-02 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.867 |  |  |  | yes | ohmic IV curve
P | 376 | 284 | 2004-02-23 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.34 |  | 0.866 |  |  |  | yes
P | 281 | 301 | 2003-12-19 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.431 |  | 0.865 |  |  |  | yes
P | 186 | 058 | 2003-12-23 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 1.37 | 1.14 | 0.84 |  |  |  | yes
P | 384 | 356 | 2004-02-25 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 2 | 2 |  | 11.243 | 0.607 | 0.766 |  |  |  | yes
P | 115 | 983 | 2004-01-06 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.626 |  | 0.738 |  |  |  | yes | IV ~linear, but <1ua
P | 323 | 212 | 2004-01-22 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 2 |  | 0.68 |  | 0.699 |  |  |  | yes | slight MD, w 8sec wait
P | 59 | 077 | 2003-12-22 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.356 |  | 0.65 |  |  |  | yes
P | 120 | 113 | 2004-01-06 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 |  | 0.636 |  |  |  | yes | slight MD, pass on 8 sec
P | 250 | 238 | 2003-11-26 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 4 |  |  | 9.5 | 1.85 | 0.628 |  |  |  | yes | 1.8 ua on 1 s wait
P | 386 | 330 | 2004-02-25 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 4 |  |  | 0.52 |  | 0.625 |  |  |  | yes
P | 392 | 374 | 2004-03-03 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.482 |  | 0.602 |  |  |  | yes | 1 abnormal sensor, afterbonding IV slightly irregular
P | 261 | 298 | 2003-12-08 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.485 |  | 0.579 |  |  |  | yes | 0.1 ua MD
P | 358 | 354 | 2004-02-10 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.59 | 0.398 | 0.553 |  |  |  | yes | 8 sec wait
P | 35 | 045 | 2003-12-08 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.563 |  | 0.526 |  |  |  | yes
P | 68 | 084 | 2003-12-26 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.364 |  | 0.422 |  |  |  | yes
P | 193 | 046 | 2003-12-22 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 0.472 |  | 0.398 |  |  |  | yes
P | 134 | 182 | 2004-01-12 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.397 |  |  |  | yes | 8 sec wait killed most MD
P | 149 | 137 | 2004-01-12 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.393 |  | 0.373 |  |  |  | yes | 8 sec IV killed most MD
P | 327 | 170 | 2004-01-27 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.347 |  | 0.371 |  |  |  | yes | slight MD, w 8sec wait
P | 157 | 106 | 2004-01-09 | PASS | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.347 |  |  |  | yes | 2 sensors w abnormal IV
P | 107 | 096 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.455 |  | 0.345 | 0.367 | NA | 0.312 | yes | 12 hr test-normal
P | 292 | 325 | 2004-01-15 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.764 |  | 0.977 |  |  |  | yes
P | 290 | 324 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.872 |  | 0.957 |  |  |  | yes
P | 285 | 319 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.965 |  | 0.925 |  |  |  | yes
P | 284 | 317 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 0.9 |  | 0.921 |  |  |  | yes
P | 8 | 014 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | No data |  | 0.917 |  |  |  | yes
P | 322 | 189 | 2004-01-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 2 |  | 0.712 |  | 0.797 |  |  |  | yes
P | 332 | 177 | 2004-01-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.677 | 0.634 | 0.778 |  |  |  | yes
P | 24 | 025 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.423 |  | 0.755 |  |  |  | yes
P | 11 | 018 | 2003-01-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.373 |  | 0.738 |  |  |  | yes | not retested, at CERN
P | 57 | 080 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.409 |  | 0.72 |  |  |  | yes
P | 54 | 075 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 |  | 0.716 |  |  |  | yes
P | 23 | 041 | 2003-02-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.333 |  | 0.706 |  |  |  | yes | not retested, at CERN
P | 10 | 017 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.332 |  | 0.672 |  |  |  | yes
P | 165 | 203 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.399 |  | 0.648 |  |  |  | yes | ohmic, no MD
P | 258 | 305 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.612 |  | 0.64 |  |  |  | yes | sensor lkge high
P | 15 | 020 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.416 |  | 0.639 |  |  |  | yes
P | 305 | 248 | 2004-01-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 0.569 |  | 0.614 |  |  |  | yes
P | 395 | 395 | 2004-03-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.596 | 0.481 | 0.608 |  |  |  | yes
P | 232 | 156 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.645 |  | 0.598 |  |  |  | yes
P | 151 | 068 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.414 |  | 0.596 |  |  |  | yes
P | 303 | 098 | 2004-01-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.577 |  | 0.595 |  |  |  | yes
P | 52 | 057 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.345 |  | 0.592 |  |  |  | yes
P | 260 | 292 | 2004-01-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.565 |  | 0.592 |  |  |  | yes
P | 255 | 289 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.589 |  | 0.586 |  |  |  | yes | early IV 21 ua sens 4
P | 310 | 291 | 2004-01-14 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.445 |  | 0.573 |  |  |  | yes
P | 307 | 296 | 2004-01-15 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.425 |  | 0.571 |  |  |  | yes
P | 312 | 276 | 2004-01-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 1 |  | 0.545 |  | 0.568 |  |  |  | yes
P | 304 | 117 | 2004-01-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.533 |  | 0.567 |  |  |  | yes
P | 318 | 247 | 2004-01-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 1 |  | 0.648 | 0.528 | 0.565 |  |  |  | yes
P | 230 | 173 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.665 |  | 0.561 |  |  |  | yes
P | 14 | 015 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.46 |  | 0.558 |  |  |  | yes
P | 256 | 288 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.613 |  | 0.558 |  |  |  | yes
P | 411 | 420 | 2004-03-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.465 |  | 0.549 |  |  |  | yes
P | 275 | 321 | 2003-12-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.487 |  | 0.539 |  |  |  | yes
P | 404 | 414 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.468 |  | 0.539 |  |  |  | yes
P | 267 | 277 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.485 |  | 0.537 |  |  |  | yes
P | 147 | 116 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.516 |  | 0.534 |  |  |  | yes
P | 409 | 418 | 2004-03-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.432 |  | 0.532 |  |  |  | yes
P | 333 | 215 | 2004-01-27 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.47 |  | 0.525 |  |  |  | yes
P | 302 | 100 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.491 |  | 0.523 |  |  |  | yes
P | 173 | 199 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.4 |  | 0.522 |  |  |  | yes
P | 50 | 072 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.523 |  | 0.52 |  |  |  | yes
P | 309 | 295 | 2004-01-14 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.461 |  | 0.516 |  |  |  | yes
P | 337 | 163 | 2004-01-28 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.497 | 0.491 | 0.513 |  |  |  | yes
P | 370 | 332 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.459 |  | 0.512 |  |  |  | yes
P | 7 | 012 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.402 |  | 0.511 |  |  |  | yes
P | 277 | 311 | 2003-12-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.547 |  | 0.51 |  |  |  | yes
P | 383 | 373 | 2004-02-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.485 |  | 0.51 |  |  |  | yes
P | 212 | 262 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.409 |  | 0.509 |  |  |  | yes
P | 47 | 065 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.455 |  | 0.507 |  |  |  | yes
P | 397 | 402 | 2004-03-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.47 |  | 0.506 |  |  |  | yes
P | 349 | 171 | 2004-02-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.471 |  | 0.504 |  |  |  | yes
P | 31 | 043 | 2003-03-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.373 |  | 0.5 |  |  |  | yes | not retested, at CERN
P | 272 | 280 | 2003-12-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.546 |  | 0.499 |  |  |  | yes
P | 215 | 217 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.642 |  | 0.497 |  |  |  | yes
P | 407 | 419 | 2004-03-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.449 |  | 0.496 |  |  |  | yes
P | 73 | 092 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.457 |  | 0.491 |  |  |  | yes
P | 306 | 294 | 2004-01-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.521 |  | 0.489 |  |  |  | yes
P | 406 | 413 | 2004-03-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.427 |  | 0.489 |  |  |  | yes
P | 146 | 135 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.54 |  | 0.488 |  |  |  | yes
P | 270 | 310 | 2003-12-15 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.504 |  | 0.487 |  |  |  | yes
P | 399 | 399 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.438 |  | 0.486 |  |  |  | yes
P | 34 | 069 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.558 |  | 0.485 |  |  |  | yes
P | 145 | 134 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.386 |  | 0.485 |  |  |  | yes
P | 246 | 256 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.465 |  | 0.484 |  |  |  | yes
P | 55 | 074 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.404 |  | 0.482 |  |  |  | yes
P | 166 | 204 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.37 |  | 0.482 |  |  |  | yes
P | 248 | 245 | 2003-11-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.456 |  | 0.482 |  |  |  | yes
P | 269 | 271 | 2003-12-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.506 |  | 0.479 |  |  |  | yes
P | 335 | 165 | 2004-01-27 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.389 |  | 0.477 |  |  |  | yes
P | 252 | 287 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.447 |  | 0.477 |  |  |  | yes
P | 84 | 094 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.67 |  | 0.475 |  |  |  | yes
P | 251 | 243 | 2003-11-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.462 |  | 0.475 |  |  |  | yes
P | 366 | 378 | 2004-02-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.92 | 0.368 | 0.475 |  |  |  | yes | possibly erroneous hi meas
P | 143 | 056 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.364 |  | 0.474 |  |  |  | yes
P | 218 | 244 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.446 |  | 0.474 |  |  |  | yes
P | 268 | 272 | 2003-12-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.518 |  | 0.473 |  |  |  | yes
P | 308 | 308 | 2004-01-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.427 |  | 0.472 |  |  |  | yes
P | 400 | 400 | 2004-03-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.438 |  | 0.472 |  |  |  | yes
P | 405 | 415 | 2004-03-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.434 |  | 0.472 |  |  |  | yes
P | 334 | 329 | 2004-01-28 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.459 |  | 0.47 |  |  |  | yes
P | 396 | 339 | 2004-03-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.457 |  | 0.468 |  |  |  | yes
P | 253 | 099 | 2003-12-02 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.44 |  | 0.467 |  |  |  | yes
P | 329 | 246 | 2004-01-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.414 |  | 0.466 |  |  |  | yes
P | 214 | 213 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.63 |  | 0.465 |  |  |  | yes
P | 315 | 309 | 2004-01-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.406 |  | 0.465 |  |  |  | yes
P | 365 | 336 | 2004-02-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.428 |  | 0.465 |  |  |  | yes
P | 381 | 362 | 2004-02-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.468 |  | 0.465 |  |  |  | yes
P | 16 | 037 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.349 |  | 0.463 |  |  |  | yes
P | 216 | 234 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.642 |  | 0.458 |  |  |  | yes
P | 247 | 257 | 2003-11-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.451 |  | 0.458 |  |  |  | yes
P | 331 | 064 | 2004-01-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.446 |  | 0.457 |  |  |  | yes
P | 239 | 266 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.42 |  | 0.455 |  |  |  | yes
P | 249 | 176 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.463 |  | 0.453 |  |  |  | yes
P | 311 | 312 | 2004-01-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.42 |  | 0.453 |  |  |  | yes
P | 276 | 313 | 2003-12-17 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.506 |  | 0.452 |  |  |  | yes
P | 220 | 239 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.544 |  | 0.451 |  |  |  | yes
P | 72 | 085 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.513 |  | 0.449 |  |  |  | yes
P | 148 | 133 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.407 |  | 0.449 |  |  |  | yes
P | 79 | 073 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.486 |  | 0.447 |  |  |  | yes
P | 213 | 255 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.492 |  | 0.447 |  |  |  | yes
P | 40 | 034 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.511 |  | 0.445 |  |  |  | yes
P | 75 | 086 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.53 |  | 0.443 |  |  |  | yes | 8 sec wait needed
P | 38 | 033 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.606 |  | 0.442 |  |  |  | yes
P | 46 | 062 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.462 |  | 0.442 |  |  |  | yes
P | 280 | 315 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.448 |  | 0.44 |  |  |  | yes
P | 94 | 118 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.501 |  | 0.439 |  |  |  | yes
P | 132 | 119 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.462 |  | 0.438 |  |  |  | yes
P | 96 | 139 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.576 |  | 0.438 |  |  |  | yes
P | 217 | 236 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.496 |  | 0.437 |  |  |  | yes
P | 177 | 180 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.344 |  | 0.433 |  |  |  | yes
P | 328 | 185 | 2004-01-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.38 |  | 0.433 |  |  |  | yes | Bad meas sens 4, extrap sum
P | 353 | 333 | 2004-02-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.398 |  | 0.433 |  |  |  | yes
P | 354 | 297 | 2004-02-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.4 |  | 0.432 |  |  |  | yes
P | 182 | 022 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.37 |  | 0.43 |  |  |  | yes
P | 90 | 153 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.428 |  | 0.429 |  |  |  | yes
P | 98 | 087 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.517 |  | 0.427 |  |  |  | yes
P | 402 | 408 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.374 |  | 0.427 |  |  |  | yes
P | 236 | 267 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.489 |  | 0.424 |  |  |  | yes
P | 142 | 112 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.496 |  | 0.422 |  |  |  | yes | 8 sec IV killed MD
P | 245 | 145 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.394 |  | 0.422 |  |  |  | yes
P | 313 | 269 | 2001-01-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.43 |  | 0.422 |  |  |  | yes | one sensor w unusual IV
P | 49 | 013 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.447 |  | 0.42 |  |  |  | yes
P | 99 | 097 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.503 |  | 0.42 |  |  |  | yes
P | 95 | 123 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.569 |  | 0.42 |  |  |  | yes
P | 367 | 251 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.363 |  | 0.42 |  |  |  | yes
P | 133 | 184 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.477 |  | 0.419 |  |  |  | yes
P | 368 | 366 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.367 |  | 0.418 |  |  |  | yes
P | 53 | 071 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.372 |  | 0.417 |  |  |  | yes
P | 169 | 196 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.46 |  | 0.417 |  |  |  | yes
P | 363 | 337 | 2004-02-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.381 |  | 0.417 |  |  |  | yes
P | 235 | 259 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.475 |  | 0.416 |  |  |  | yes
P | 228 | 263 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.51 |  | 0.415 |  |  |  | yes
P | 227 | 249 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.552 |  | 0.414 |  |  |  | yes
P | 77 | 088 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.495 |  | 0.413 |  |  |  | yes
P | 51 | 026 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.485 |  | 0.412 |  |  |  | yes
P | 362 | 340 | 2004-02-13 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.35 |  | 0.412 |  |  |  | yes
P | 138 | 154 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.408 |  | 0.411 |  |  |  | yes
P | 350 | 357 | 2004-02-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.376 |  | 0.411 |  |  |  | yes
P | 347 | 383 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.368 |  | 0.411 |  |  |  | yes
P | 130 | 164 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.482 |  | 0.406 |  |  |  | yes
P | 326 | 148 | 2004-01-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.409 |  | 0.403 |  |  |  | yes
P | 244 | 155 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.38 |  | 0.403 |  |  |  | yes
P | 373 | 363 | 2004-02-19 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  | 1 |  | 0.4 |  | 0.403 |  |  |  | yes
P | 87 | 093 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.489 |  | 0.402 |  |  |  | yes
P | 185 | 211 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.448 | 0.393 | 0.401 |  |  |  | yes
P | 167 | 059 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.405 |  | 0.4 |  |  |  | yes
P | 65 | 089 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.438 |  | 0.399 |  |  |  | yes
P | 135 | 103 | 2004-01-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.425 |  | 0.399 |  |  |  | yes
P | 175 | 108 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.385 |  | 0.399 |  |  |  | yes
P | 170 | 198 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.388 |  | 0.399 |  |  |  | yes
P | 183 | 229 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.391 |  | 0.399 |  |  |  | yes
P | 33 | 044 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.46 |  | 0.398 |  |  |  | yes
P | 282 | 304 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.375 |  | 0.398 |  |  |  | yes
P | 274 | 306 | 2003-12-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.482 |  | 0.398 |  |  |  | yes
P | 117 | 179 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.66 |  | 0.396 |  |  |  | yes
P | 369 | 342 | 2004-02-18 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.344 |  | 0.396 |  |  |  | yes
P | 48 | 070 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.435 |  | 0.395 |  |  |  | yes
P | 226 | 254 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.528 |  | 0.395 |  |  |  | yes
P | 238 | 270 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.398 |  | 0.394 |  |  |  | yes | 8sec IV killed MD
P | 371 | 331 | 2004-02-19 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.362 |  | 0.394 |  |  |  | yes
P | 374 | 384 | 2004-02-20 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.402 |  | 0.394 |  |  |  | yes
P | 210 | 258 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.491 |  | 0.393 |  |  |  | yes
P | 208 | 224 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.525 |  | 0.392 |  |  |  | yes
P | 17 | 038 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.384 |  | 0.391 |  |  |  | yes
P | 92 | 110 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.469 |  | 0.391 |  |  |  | yes
P | 352 | 334 | 2004-02-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.369 |  | 0.391 |  |  |  | yes
P | 80 | 159 | 2003-12-26 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.468 |  | 0.389 |  |  |  | yes
P | 211 | 250 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.469 |  | 0.389 |  |  |  | yes
P | 70 | 090 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.492 |  | 0.388 |  |  |  | yes
P | 225 | 241 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.504 |  | 0.388 |  |  |  | yes
P | 382 | 380 | 2004-02-24 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.388 |  | 0.388 |  |  |  | yes
P | 18 | 028 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.397 |  | 0.387 |  |  |  | yes
P | 336 | 109 | 2004-01-27 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.325 |  | 0.387 |  |  |  | yes
P | 64 | 083 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.403 |  | 0.386 |  |  |  | yes
P | 153 | 187 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.38 |  | 0.386 |  |  |  | yes
P | 206 | 237 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.522 |  | 0.386 |  |  |  | yes
P | 359 | 353 | 2004-02-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.348 |  | 0.385 |  |  |  | yes
P | 361 | 361 | 2004-02-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.358 |  | 0.385 |  |  |  | yes
P | 351 | 349 | 2004-02-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.382 |  | 0.384 |  |  |  | yes
P | 372 | 385 | 2004-02-19 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.354 |  | 0.382 |  |  |  | yes
P | 128 | 101 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.444 |  | 0.38 |  |  |  | yes
P | 403 | 416 | 2004-03-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.392 |  | 0.379 |  |  |  | yes
P | 164 | 202 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.331 |  | 0.378 |  |  |  | yes
P | 41 | 053 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.435 |  | 0.375 |  |  |  | yes
P | 364 | 381 | 2004-02-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.359 |  | 0.374 |  |  |  | yes
P | 42 | 036 | 2003-12-08 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.412 |  | 0.373 |  |  |  | yes
P | 341 | 293 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.358 |  | 0.373 |  |  |  | yes
P | 113 | 067 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.62 |  | 0.372 |  |  |  | yes
P | 314 | 107 | 2004-01-16 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.326 |  | 0.371 |  |  |  | yes
P | 343 | 264 | 2004-01-30 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.339 |  | 0.371 |  |  |  | yes
P | 60 | 078 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.425 |  | 0.37 |  |  |  | yes
P | 265 | 303 | 2003-12-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.339 |  | 0.37 |  |  |  | yes
P | 348 | 365 | 2004-02-05 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.382 |  | 0.37 |  |  |  | yes
P | 43 | 048 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.468 |  | 0.368 |  |  |  | yes
P | 129 | 052 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.439 |  | 0.365 |  |  |  | yes
P | 104 | 151 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.442 |  | 0.364 |  |  |  | yes
P | 171 | 201 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.348 |  | 0.361 |  |  |  | yes
P | 340 | 328 | 2004-01-28 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.377 |  | 0.361 |  |  |  | yes
P | 105 | 152 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.406 |  | 0.358 |  |  |  | yes
P | 360 | 382 | 2004-02-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.321 |  | 0.358 |  |  |  | yes
P | 172 | 207 | 2003-12-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.435 |  | 0.357 |  |  |  | yes
P | 221 | 242 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.393 |  | 0.357 |  |  |  | yes
P | 240 | 233 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.379 |  | 0.356 |  |  |  | yes
P | 19 | 040 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.386 |  | 0.355 |  |  |  | yes | 8 sec wait eliminated MD
P | 205 | 235 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.431 |  | 0.354 |  |  |  | yes
P | 346 | 379 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.328 |  | 0.354 |  |  |  | yes
P | 288 | 221 | 2004-01-07 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.346 |  | 0.353 |  |  |  | yes
P | 342 | 314 | 2004-01-30 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.348 |  | 0.353 |  |  |  | yes
P | 32 | 027 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.368 |  | 0.352 |  |  |  | yes
P | 181 | 226 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.369 |  | 0.352 |  |  |  | yes
P | 61 | 082 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.339 |  | 0.351 |  |  |  | yes
P | 27 | 042 | 2003-12-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.345 |  | 0.35 |  |  |  | yes
P | 178 | 208 | 2003-12-11 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.325 |  | 0.35 |  |  |  | yes
P | 264 | 252 | 2003-12-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.347 |  | 0.35 |  |  |  | yes
P | 144 | 130 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.417 |  | 0.349 |  |  |  | yes
P | 203 | 220 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 | 1 |  |  | 0.459 |  | 0.348 |  |  |  | yes
P | 379 | 338 | 2004-02-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.333 |  | 0.345 |  |  |  | yes
P | 345 | 377 | 2004-02-04 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.31 |  | 0.343 |  |  |  | yes
P | 263 | 160 | 2003-12-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.343 |  | 0.342 |  |  |  | yes
P | 156 | 122 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.383 |  | 0.338 |  |  |  | yes
P | 121 | 993 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.446 |  | 0.337 |  |  |  | yes
P | 123 | 105 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.448 |  | 0.336 |  |  |  | yes
P | 266 | 253 | 2003-12-10 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.32 |  | 0.336 |  |  |  | yes
P | 150 | 136 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.368 |  | 0.334 |  |  |  | yes
P | 122 | 980 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.463 |  | 0.332 |  |  |  | yes | 8 sec
P | 102 | 132 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.471 |  | 0.331 |  |  |  | yes
P | 158 | 191 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.353 |  | 0.331 |  |  |  | yes
P | 155 | 138 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 | 1 |  |  | 0.372 |  | 0.33 |  |  |  | yes
P | 154 | 126 | 2004-01-12 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.359 |  | 0.329 |  |  |  | yes
P | 375 | 231 | 2004-02-20 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.347 |  | 0.328 |  |  |  | yes
P | 101 | 121 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | * | 1 | 0 |  |  |  | 0.419 |  | 0.323 |  |  |  | yes
P | 159 | 188 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.371 |  | 0.323 |  |  |  | yes
P | 152 | 141 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.346 |  | 0.322 |  |  |  | yes
P | 344 | 375 | 2004-02-03 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.284 |  | 0.322 |  |  |  | yes
P | 199 | 218 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.377 |  | 0.32 |  |  |  | yes
P | 112 | 111 | 2004-01-06 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.523 |  | 0.318 |  |  |  | yes
P | 187 | 230 | 2003-12-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.398 |  | 0.316 |  |  |  | yes
P | 161 | 193 | 2003-11-25 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.317 |  | 0.315 |  |  |  | yes
P | 160 | 131 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.31 |  | 0.313 |  |  |  | yes
P | 196 | 227 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.37 |  | 0.312 |  |  |  | yes
P | 197 | 166 | 2003-12-21 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.341 |  | 0.31 |  |  |  | yes
P | 162 | 161 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.305 |  | 0.309 |  |  |  | yes
P | 377 | 335 | 2004-02-23 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.298 |  | 0.309 |  |  |  | yes
P | 195 | 206 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.308 |  | 0.3 |  |  |  | yes
P | 194 | 054 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | 0.317 |  | 0.293 |  |  |  | yes
P | 163 | 169 | 2004-01-09 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | ** | 0 | 1 |  |  |  | inc data |  | 0.285 |  |  |  | yes | missing sensor data, OK
P | 67 | 095 | 2003-12-22 | PASS | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.312 |  | 0.257 |  |  |  | yes | 8 sec wait killed MD
406

## Units on HOLD for IV reasons
| Module ID | [18 empty] | I(500) |  |  | MDM (Microdischarge Decay Measurement)
| P# | Hybrid | Date tested | PASS/ HOLD | I-V Score | [7 empty] | Gold |  |  | Quality CLASS |  |  | Sum 4 sens | Sum 4 sens | Module | Module | Time to baseline | Base current | Plotted | Comment
[13 empty] | Stars |  |  | 2-3 | 4-5 | 6+ |  | (after MDM) | afterbonding | Post-MDM
MODULES
P | 28 | 030 | 2003-12-04 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.378 |  | 21.3 | 2.3 | >12 hrs | 2.5 | yes | Very noisy
P | 63 | 081 | 2003-12-22 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.355 |  | 21 | 21 | >12 hrs | 21 ua | yes | 21 ua at 200 V
P | 241 | 268 | 2003-12-01 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.467 | 0.385 | 21.4 | 4.99 | >12 hrs | 4.8 | yes | "base" 4.8 ua in 2 hrs
P | 391 | 359 | 2004-02-27 | HOLD | 5 | 0 | 0 | 0 | 0 | 0 | 1 | 0 |  | 0 | 0 |  |  |  | 0.411 |  | 21.4 | 21.1 | >12 hrs | 21 | yes | 21 ua at 200V, at 450 V after 12 hrs
4-WAFER ASSEMBLIES
P | 25 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 98.3 | [5 empty] | yes | One really bad sensor
P | 126 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | >300 | [5 empty] | yes
P | 174 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 |  |  | 5.036 | 0.931 |  |  |  |  | yes | One bad sensor, seems to condition
P | 192 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 385 | [5 empty] | yes | One insanely high sensor
P | 223 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 347 | 0.445 |  |  |  |  | yes | One chipped sensor, remeas next day was OK??
P | 293 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 4 |  | 23.238 | 17.531 |  |  |  |  | yes | 3 abnormal sensors, Nobu suggested also to hold
P | 298 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 3 |  | 39.453 | 21.975 |  |  |  |  | yes | One very bad sensor, Nobu suggested holding
P | 320 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  | 2 |  | 2.782 | 2.77 |  |  |  |  | yes | One bad sensor, no conditioning effect
P | 388 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 | 1 | 2 | 1 | 48.378 | 26.413 | [5 empty] | May possibly condition
P | 394 |  | HOLD | hold-IV |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 0.566 | 0.493 | [5 empty] | Very noisy, NOT probe-related
P | 451 |  | IV & glue | hold iv |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.854 | 2.91 | [5 empty] | One very abnormal sensor
P | 458 |  |  | HOLD |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 |  | 0 | 0 |  |  |  | 21.96 | 21.91 | [5 empty] | 1 sensor, 21ua @ 30V no condition effect