|  | Thermal Qualification of ATLAS Pixel Detector Disk Sectors |
| --- | --- |
| ATLAS Project Document No: | Institute Document No. | Created : 6/29/01 | Page: 1 of 2 |
|  |  | Modified : | Rev. No .: |
| Thermal Qualification of ATLAS Pixel Detector Disk Sectors Abstract The thermal qualification tests for the ATLAS Pixel Detector disk sectors are described. Recent measurements on prototypes are presented. |
| Prepared by: M. G. D. Gilchriese, T. Johnson, F. McCormack, T. Weber, J. Wirth, W. Wise | Checked by: | Approved by: |
| Distribution List |

| History of Changes |
| --- |
| Rev. No. | Date | Pages | Description of changes |

Table of Contents

1	Introduction	4

2	Thermal Qualification	4

2.1	Qualification Tests	4

2.2	Test Procedures	5

Dummy heaters (platinum-coated silicon chips) were used in place of actual pixel modules. Each of the heaters had a resistance of 6-7 ohms: the three on each side were connected in series to a power supply at around 20 V to provide about 25 W of power to each side of the sector. The coolant used in the test is water at around 21o C. The flow rate in all tests was around .75 liters/minute. Coolant temperature and flow varied very little between tests.	5

3	Test Results	5

3.1	Baseline Testing	5

3.2	Temperature cycling test	6

3.3	Cold Shock	7

3.4	Pressure Testing	7

3.5	Thermal cycling	8

3.6	Radiation Testing	8

3.7	Thermal cycling	8

3.8	Thermal Shock	9

3.9	Pressure Testing	9

3.10	Loss of coolant test	9

4	Conclusions	9

5	Heading 1	9

5.1	Heading 2	9

5.1.1	Heading 3	9

# Introduction

This note summarizes recent thermal qualification tests of the pixel detector disk sectors. The disk sectors support pixel modules in the disk (end) regions of the pixel detector. The components of a disk sector are shown in Figure 1.

Figure 1. A disk sector without modules.

# Thermal Qualification Criteria and Test Procedures

## Thermal Qualification Criteria

The thermal requirements, including stress testing, are defined in the document on ATLAS Pixel Local Supports Requirements. The thermal qualification tests used to qualify the sector design are summarized in Table 1.

50 thermal cycles 20o C to -35o C

Thermal shock 20o C to -35o C in < 60 sec

Pressure to 8 bar absolute

Irradiate to 50 Mrad

50 thermal cycles 20o C to -35o C

Thermal shock 20o C to -35o C in < 60 sec

Pressure to 8 bar absolute

Loss of coolant (sector in air at 20o C)

Table 1. The thermal qualification test sequence.

## Test Procedures

Dummy heaters were used in place of actual pixel modules. The dummy heaters are platinum-coated silicon. The thin platinum is deposited by evaporation on the silicon in a footprint that simulates the heat distribution expected for the pixel module electronics. Each of the heaters had a resistance of 6-7 ohms. The three on each side were connected in series to a power supply at around 20 V to provide about 25 W of power to each side of the sector. The coolant used in the test is water at around 21o C. The flow rate in all tests was around 0.75 liters/minute Prior to testing, the sectors and heaters were coated with black paint to ensure uniform surface emissivity. Thermal cycling damaged the paint in some small areas, but no important data were lost. Tests were done in air at room temperature. Unpainted heaters on a sector are shown in Figure 2.

Figure 2. Unpainted dummy heaters on a disk sector.

****Describe IR camera and so forth here***. The temperature difference between the incoming coolant and the hottest point on the sector is an important measure of sector performance. This value, along with the thermograms after each test, are given in the subsequent sections.

# Prototypes

The test results on two prototypes are described in this note. The two prototypes differ only in the thermal compound used to connect the aluminum tube to the sector facings. Prototype 11 used CGL-7018, a thermally conductive paste adhesive, between the aluminum tube and the sector facings. Prototype 12 used**** an epoxy-based, thermally conducting adhesive.

# Test Results

## Baseline Testing

The baseline heat pattern for sector 11 is shown in Figure 3, and the heat pattern for sector 12 is shown in Figure 4. The test results are summarized in Table 2.

Figure 3. Sector 11 baseline.                                               Figure 4. Sector 12 baseline.

| Conditions (average) | Sector 11 | Sector 12 |
| --- | --- | --- |
| Coolant temperature (o C) | 21.0 | 21.0 |
| Coolant flow (liters/min) | 0.77 | 0.75 |
| Power to heaters (W) | 48.4 | 49.1 |
| Maximum (T (o C) | 6.9 | 7.8 |

Table 2. Baseline test results.

## Temperature cycling test

The sector was cycled 50 times between 20o C and -35o C. The results are shown in

Figure 5. Sector 11, after thermal cycling.                          Figure 6. Sector 12, after thermal cycling.

| Conditions (average) | Sector 11 | Sector 12 |
| --- | --- | --- |
| Coolant temperature (o C) | 21.0 | 21.0 |
| Coolant flow (liters/min) | 0.77 | 0.75 |
| Power to heaters (W) | 48.4 | 49.1 |
| DT (o C) | 7.7 | 7.4 |

Table 3. Results after temperature cycling.

Figure 5, Figure 6 and Table 3.

## Cold Shock

Sector 11 was brought from room temperature at 17.3o C to –37o C in 45 seconds by pumping nitrogen vapor into the coolant tube. The average cooling rate was 1.2o C/sec. The flow of nitrogen was then shut off, and the sample was allowed to warm to room temperature. Sector 12 underwent similar treatment, being cooled from 20o C to -40o C in 50 seconds. The average ramp rate was 1.2o C/sec. The sectors were then tested in normal operation. Thermal images are given below in Figure 7 and Figure 8. Damage to the paint on the middle heater of sector 11 caused some minor loss of data. Results are summarized in Table 4.

Figure 7. Sector 11, after cold shock.                            Figure 8. Sector 12, after cold shock.

| Conditions (average) | Sector 11 | Sector 12 |
| --- | --- | --- |
| Coolant temperature (o C) | 21.0 | 21.0 |
| Coolant flow (liters/min) | 0.77 | 0.75 |
| Power to heaters (W) | 48.4 | 49.1 |
| Maximum (T (o C) | 7.6 | 7.5 |

Table 4. Results after thermal shock.

## Pressure Testing

The coolant tube was attached to a compressed gas cylinder and subjected to 8 bar for 1 hour. The results are given below.

Figure 9. Sector 11, after pressure test.                              Figure 10. Sector 12, after pressure test.

| Conditions (average) | Sector 11 | Sector 12 |
| --- | --- | --- |
| Coolant temperature (o C) | 21.0 | 20.8 |
| Coolant flow (liters/min) | 0.76 | 0.76 |
| Power to heaters (W) | 48.4 | 49.1 |
| Maximum (T (o C) | 7.6 | 7.0 |

Table 5. Results after pressure test.

## Additional Thermal Cycling

The sectors were exposed to an additional 50 cycles from 20o C to -35o C, for a total of 100 cycles. The cycling damaged paint on sector 12. However, most of the sector was unaffected, and accurate temperature readings were possible. The results are given below.

Figure 11. Sector 11, after additional thermal cycling.       Figure 12. Sector 12, after additional thermal cycling.

| Conditions (average) | Sector 11 | Sector 12 |
| --- | --- | --- |
| Coolant temperature (o C) | 21.0 | 20.8 |
| Coolant flow (liters/min) | 0.76 | 0.75 |
| Power to heaters (W) | 48.4 | 49.1 |
| Maximum (T (o C) | 7.6 | 7.2 |

Table 6. Results after additional thermal cycling.

## Radiation Testing

## Thermal cycling

## Thermal Shock

## Pressure Testing

## Loss of coolant test

# Conclusions

All results are summarized in

|  | Maximum ( T for sector 11 (o C) | Maximum ( T for sector 12 (o C) |
| --- | --- | --- |
| Baseline | 6.9 | 7.8 |
| 50 T cycles 20o C to -35o C | 7.7 | 7.4 |
| Thermal shock | 7.6 | 7.5 |
| Pressure | 7.6 | 7.0 |
| 50 more T cycles 20o C to -35o C | 7.1 | 7.3 |

Table 7. Summary of all results on thermal qualification.

# Comments

Enter No

Enter No

Enter Date dd/MM/yyy

Enter Date dd/MM/yyy

Enter No

Enter text of abstract

Enter Name, Institute, E-mail

Enter Number/Title

# Endnotes

ATL-IP-EP-0005. “ATLAS Pixel Local Supports Requirements.” June 2001.

AI Technologies CGL-7018-LV-XF.